## Sheet1
Unshielded reference bridge. Rvb = 2K2 CF					Bridge evaluation by differential method. Version 1.00. DWK Oct. 2007.
		+/- esd					Cells with yellow background are input data.
R0cal =	49.90	0.06	<< Actual ref. load resistance (measured).
Design R0	50.00
k =	0.96	0.01
Ri =	50.04	0.06
N =	12
C2 + Cx	5.350		pF
Cx	0.223		pF	<< zero when Faraday shield is present.
C2	5.128
dR0 / dC1	0.748		Ohms / pF
x cal	3.342	0.005	<<< to determine the final calibration setting, vary xcal to get best true accuracy.
C1a cal	18.991		pF
				Magnitude error vs f
				f	C Scale	C1a	C1a – cal		R0	delta R0	|Z0|	delta|Z0|	|deltaZ0|
				/ MHz	x	/ pF	delta C1	delta R0	R0cal+dRo	/ R0		/ R0
				1.6	3.375	18.882	-0.110	0.082	49.982	0.001645	49.982	-0.000358	0.017918
				1.7	3.375	18.882	-0.110	0.082	49.982	0.001645	49.982	-0.000358	0.017919
				1.8	3.375	18.882	-0.110	0.082	49.982	0.001645	49.982	-0.000358	0.017924
				1.9	3.375	18.882	-0.110	0.082	49.982	0.001645	49.982	-0.000358	0.017920
				2.0	3.375	18.882	-0.110	0.082	49.982	0.001645	49.982	-0.000358	0.017921
				2.2	3.380	18.865	-0.126	0.095	49.995	0.001894	49.995	-0.000110	0.005489
				2.4	3.385	18.848	-0.143	0.107	50.007	0.002143	50.007	0.000139	0.006952
				2.6	3.385	18.848	-0.143	0.107	50.007	0.002143	50.007	0.000139	0.006949
				2.8	3.385	18.848	-0.143	0.107	50.007	0.002143	50.007	0.000139	0.006948
				3.0	3.385	18.848	-0.143	0.107	50.007	0.002143	50.007	0.000139	0.006946
				3.5	3.385	18.848	-0.143	0.107	50.007	0.002143	50.007	0.000139	0.006946
				4.0	3.385	18.848	-0.143	0.107	50.007	0.002143	50.007	0.000139	0.006946
				4.5	3.385	18.848	-0.143	0.107	50.007	0.002143	50.007	0.000139	0.006946
				5.0	3.390	18.832	-0.160	0.119	50.019	0.002392	50.019	0.000388	0.019382
				6.0	3.390	18.832	-0.160	0.119	50.019	0.002392	50.019	0.000388	0.019381
				7.0	3.390	18.832	-0.160	0.119	50.019	0.002392	50.019	0.000388	0.019381
				8.0	3.390	18.832	-0.160	0.119	50.019	0.002392	50.019	0.000388	0.019381
				9.0	3.390	18.832	-0.160	0.119	50.019	0.002392	50.019	0.000388	0.019381
				10.0	3.385	18.848	-0.143	0.107	50.007	0.002143	50.007	0.000139	0.006946
				12.0	3.380	18.865	-0.126	0.095	49.995	0.001894	49.995	-0.000110	0.005489
				14.0	3.380	18.865	-0.126	0.095	49.995	0.001894	49.995	-0.000110	0.005489
				16.0	3.380	18.865	-0.126	0.095	49.995	0.001894	49.995	-0.000110	0.005488
				18.0	3.380	18.865	-0.126	0.095	49.995	0.001894	49.995	-0.000110	0.005485
				20.0	3.380	18.865	-0.126	0.095	49.995	0.001894	49.995	-0.000110	0.005482
				22.0	3.380	18.865	-0.126	0.095	49.995	0.001894	49.995	-0.000110	0.005484
				24.0	3.380	18.865	-0.126	0.095	49.995	0.001894	49.995	-0.000110	0.005485
				26.0	3.380	18.865	-0.126	0.095	49.995	0.001894	49.995	-0.000110	0.005487
				28.0	3.380	18.865	-0.126	0.095	49.995	0.001894	49.995	-0.000110	0.005490
				30.0	3.375	18.882	-0.110	0.082	49.982	0.001645	49.982	-0.000358	0.017923
			Magnitude precision of the bridge +/- >>>					0.0194	Ohms
			Magnitude precision as a percentagee +/- >>>					0.0388	%
			True (claimable) accuracy of the bridge +/- >>>					0.0655	Ohms
			True accuracy as a percentagee +/- >>>					0.1309	%
KT =	12.117		<< Transformer constant.
C1 =	62.13		pF	<< estimated C1
C1b =	38.90		pF	<< padding capacitance. Nominal will do if no measurement is available.
C1s =	4.24		pF	<< estimated contribution to C1 from strays.

## Sheet2
Unshielded reference bridge. Rvb = 2K2 CF					T = 22 deg. C
				Cells with yelow background are input data.
R0	49.90
k	0.96
Ri	50.04
N	12
C2+Cx	5.350		pF	<<< must include strays.
Rvcal	2516.70		<<< to get final calibration setting, vary this to minimise max |phase error|
Li calc	8.062		uH	<<< should be ca. 1% lower than measured sec. Inductance. If not, C2+Cx has been wrongly estimated.
Lsec	8.150		uH	<<< measured sec. Inductance. For comparison
Li/Lsec	0.989

				f	Rv	Rv-Rvcal	dX0 / dRv	X0	delta phi	| delta phi |
				/MHz	/ Ohms		Deriv				Phase error vs f
				1.6	2519.0	2.3	0.011729	-0.027	-0.0309	0.0309
				1.7	2519.0	2.3	0.011039	-0.025	-0.0291	0.0291
				1.8	2518.0	1.3	0.010434	-0.014	-0.0155	0.0155
				1.9	2519.0	2.3	0.009877	-0.023	-0.0260	0.0260
				2.0	2519.0	2.3	0.009383	-0.022	-0.0247	0.0247
				2.2	2518.0	1.3	0.008537	-0.011	-0.0127	0.0127
				2.4	2520.0	3.3	0.007813	-0.026	-0.0295	0.0295
				2.6	2519.0	2.3	0.007218	-0.017	-0.0190	0.0190
				2.8	2519.0	2.3	0.006702	-0.015	-0.0177	0.0177
				3.0	2518.0	1.3	0.006260	-0.008	-0.0093	0.0093
				3.5	2517.0	0.3	0.005370	-0.002	-0.0018	0.0018
				4.0	2516.0	-0.7	0.004703	0.003	0.0038	0.0038
				4.5	2516.0	-0.7	0.004180	0.003	0.0034	0.0034
				5.0	2515.0	-1.7	0.003765	0.006	0.0073	0.0073
				6.0	2516.0	-0.7	0.003135	0.002	0.0025	0.0025
				7.0	2516.0	-0.7	0.002687	0.002	0.0022	0.0022
				8.0	2516.0	-0.7	0.002351	0.002	0.0019	0.0019
				9.0	2518.0	1.3	0.002087	-0.003	-0.0031	0.0031
				10.0	2517.0	0.3	0.001880	-0.001	-0.0006	0.0006
				12.0	2510.0	-6.7	0.001575	0.011	0.0121	0.0121
				14.0	2509.0	-7.7	0.001351	0.010	0.0119	0.0119
				16.0	2506.0	-10.7	0.001185	0.013	0.0145	0.0145
				18.0	2496.0	-20.7	0.001062	0.022	0.0252	0.0252
				20.0	2488.0	-28.7	0.000962	0.028	0.0316	0.0316
				22.0	2489.0	-27.7	0.000874	0.024	0.0277	0.0277
				24.0	2489.0	-27.7	0.000801	0.022	0.0254	0.0254
				26.0	2496.0	-20.7	0.000735	0.015	0.0174	0.0174
				28.0	2512.0	-4.7	0.000674	0.003	0.0036	0.0036
				30.0	2542.0	25.3	0.000614	-0.016	-0.0178	0.0178
				0.0316	Max |phase error|.

## Sheet3
Data reduction for 3-30pF beehive trimmer with reference markings..
		Dial	C /pF	esd C	fit	w	wy	wx	wx^2	wxy	y calc	obs-calc	we	we^2
		x	y					C calc	e
0 – 250pF range >>		0.0	30.0	0.25	1	16.000	480.00	0.00	0.00	0.00	30.10	-0.10	-1.61	0.163
		0.5	28.5	0.25	1	16.000	456.00	8.00	4.00	228.00	28.44	0.06	0.98	0.060
		1.0	26.5	0.25	1	16.000	424.00	16.00	16.00	424.00	26.78	-0.28	-4.43	1.224
		1.5	25.0	0.25	1	16.000	400.00	24.00	36.00	600.00	25.11	-0.11	-1.83	0.210
0-25pF range >>		2.0	23.7	0.10	1	100.000	2370.00	200.00	400.00	4740.00	23.45	0.25	24.76	6.129
		2.5	21.9	0.10	1	100.000	2190.00	250.00	625.00	5475.00	21.79	0.11	10.97	1.203
		3.0	20.1	0.10	1	100.000	2010.00	300.00	900.00	6030.00	20.13	-0.03	-2.82	0.080
		3.5	18.4	0.10	1	100.000	1840.00	350.00	1225.00	6440.00	18.47	-0.07	-6.62	0.438
		4.0	16.7	0.10	1	100.000	1670.00	400.00	1600.00	6680.00	16.80	-0.10	-10.41	1.083
		4.5	15.0	0.10	1	100.000	1500.00	450.00	2025.00	6750.00	15.14	-0.14	-14.20	2.016
		5.0	13.4	0.10	1	100.000	1340.00	500.00	2500.00	6700.00	13.48	-0.08	-7.99	0.638
		5.5	11.8	0.10	1	100.000	1180.00	550.00	3025.00	6490.00	11.82	-0.02	-1.78	0.032
		6.0	10.1	0.10	1	100.000	1010.00	600.00	3600.00	6060.00	10.16	-0.06	-5.57	0.310
		6.5	8.5	0.10	1	100.000	850.00	650.00	4225.00	5525.00	8.49	0.01	0.64	0.004
		7.0	6.9	0.10	1	100.000	690.00	700.00	4900.00	4830.00	6.83	0.07	6.85	0.469
		7.5	5.3	0.10	1	100.000	530.00	750.00	5625.00	3975.00	5.17	0.13	13.06	1.705
Starting to unmesh >>		8.0	4.1	0.10	0	0.000	0.00	0.00	0.00	0.00	3.51	0.59	0.00	0.000
		8.5	3.5	0.10	0	0.000	0.00	0.00	0.00	0.00	1.85	1.60	0.00	0.000
		9.0	3.1	0.10	0	0.000	0.00	0.00	0.00	0.00	0.18	2.92	0.00	0.000
		9.5	2.9	0.10	0	0.000	0.00	0.00	0.00	0.00	-1.48	4.38	0.00	0.000
		10.0	2.8	0.10	0	0.000	0.00	0.00	0.00	0.00	-3.14	5.94	0.00	0.000
		sums >>>>>>>>>			16	1264.000	18940.00	5748.00	30706.00	70947.00			8.32E-12	15.763
D	5772880
a	30.100796	Capacitance / pF vs logging scale
esd a	0.005989
b	-3.324183
esd b	0.000247
Chi^2/n-2)	1.125893
esd fit	1.061081