## Sheet1
FT50-61, 12 turns 0.9mm wire. Lsec = 9.25uH @ 1.5915MHz. 500pF variable ref. cap. Faraday shield earthed on gen. side.					D. W. Knight. June 2007.				Version	1.00
		esd					Show columns N-Q to see deatails of fitting procedure.					See sheet 2 for C1'm vs frequency
N =	12.00					Li' = N R0 C2 Rv					Ci computed from Rv
R0 =	49.89	0.06					from fit	from fit
C2/pF	10.30	0.20					Ci =	Li =
L2 / nH	50.00	20.00	< estimated
L2 fit	1	< 1 to include L2 in fit, 0 to exclude it.
1/(N R0 C2)	1.62E+08					7.993	8.947
			Log (f)	f	x	Rv	esd Rv	Li'	Corr(L2)	y	esd y	fit	weight	w x	w x^2	w y	w x y	y calc	obs-calc	w e	w e^2	Li' calc	Li'					Rv Calc	obs-calc	Computed	Computed
				/MHz	-(2 pi f)^2	Ohms	Ohms	/ uH					w					e			/ uH	obs-calc					Li /uH	Ci /pF
			6.15	1.4	-77.378	1462.6	14.63	9.0190	-4.4	110873	1108.77	1	8.13E-07	-0.000062941	0.004870	0.090186	-6.978402983	111151	-277.9	-0.000226	0.06	8.9964	0.0225					1458.9	3.7	8.969	11.527
			6.18	1.5	-88.826	1462.6	7.31	9.0190	-5.1	110872	554.39	1	3.25E-06	-0.000289013	0.025672	0.360743	-32.043498224	111059	-187.0	-0.000609	0.11	9.0038	0.0152					1460.1	2.5	8.962	10.042
			6.20	1.6	-101.065	1462.5	0.88	9.0184	-5.8	110879	66.53	1	2.26E-04	-0.022832470	2.307558	25.049712	-2531.642896500	110961	-82.3	-0.018602	1.53	9.0117	0.0067	<< same meas freq as Lsec					1461.4	1.1	8.953	8.751
			6.23	1.7	-114.093	1464.2	0.88	9.0288	-6.5	110750	66.45	1	2.26E-04	-0.025835676	2.947660	25.078662	-2861.290395195	110857	-107.7	-0.024385	2.63	9.0201	0.0088					1462.8	1.4	8.955	8.880
			6.26	1.8	-127.910	1464.7	0.88	9.0319	-7.3	110711	66.43	1	2.27E-04	-0.028984350	3.707390	25.087047	-3208.886032605	110747	-35.8	-0.008121	0.29	9.0290	0.0029					1464.2	0.5	8.949	8.216
			6.28	1.9	-142.517	1466.9	0.88	9.0455	-8.1	110544	66.33	1	2.27E-04	-0.032391377	4.616325	25.124539	-3580.676156629	110630	-86.0	-0.019536	1.68	9.0385	0.0070					1465.8	1.1	8.953	8.539
			6.30	2.0	-157.914	1467.1	0.88	9.0467	-9.0	110528	66.32	1	2.27E-04	-0.035900509	5.669181	25.127765	-3968.017678655	110507	21.2	0.004812	0.10	9.0485	-0.0017					1467.4	-0.3	8.945	7.802
			6.34	2.2	-191.076	1469.5	0.88	9.0615	-10.9	110346	66.21	1	2.28E-04	-0.043581857	8.327427	25.168442	-4809.073606382	110242	103.8	0.023682	2.46	9.0701	-0.0085					1470.9	-1.4	8.938	7.393
			6.38	2.4	-227.396	1473.0	0.88	9.0831	-12.9	110082	66.06	1	2.29E-04	-0.052113370	11.850356	25.227915	-5736.719009859	109952	129.9	0.029766	3.87	9.0938	-0.0107					1474.7	-1.7	8.936	7.365
			6.41	2.6	-266.874	1478.2	0.89	9.1152	-15.1	109692	65.82	1	2.31E-04	-0.061593414	16.437687	25.316460	-6756.307554660	109636	56.0	0.012916	0.72	9.1198	-0.0047					1479.0	-0.8	8.941	7.727
			6.45	2.8	-309.511	1483.2	0.89	9.1460	-17.4	109320	65.60	1	2.32E-04	-0.071917848	22.259350	25.401535	-7862.049229725	109295	24.6	0.005711	0.14	9.1481	-0.0021					1483.5	-0.3	8.944	7.857
			6.48	3.0	-355.306	1488.7	0.89	9.1799	-19.9	108913	65.36	1	2.34E-04	-0.083172177	29.551554	25.495127	-9058.565582665	108929	-15.8	-0.003707	0.06	9.1786	0.0013					1488.5	0.2	8.947	7.981
			6.54	3.5	-483.611	1501.9	0.90	9.2613	-26.9	107949	64.79	1	2.38E-04	-0.115223035	55.723083	25.719488	-12438.217248158	107904	45.3	0.010805	0.49	9.2652	-0.0039					1502.5	-0.6	8.941	7.844
			6.60	4.0	-631.655	1518.7	0.91	9.3649	-34.7	106747	64.07	1	2.44E-04	-0.153881055	97.199689	26.005198	-16426.305146592	106720	26.4	0.006428	0.17	9.3672	-0.0023					1519.1	-0.4	8.942	7.896
			6.65	4.5	-799.438	1538.9	0.92	9.4895	-43.4	105336	63.23	1	2.50E-04	-0.199970997	159.864405	26.348812	-21064.240147936	105379	-42.8	-0.010708	0.46	9.4856	0.0039					1538.3	0.6	8.947	7.992
			6.70	5.0	-986.960	1560.8	0.94	9.6245	-52.8	103848	62.34	1	2.57E-04	-0.253954380	250.642926	26.721198	-26372.765573915	103880	-32.0	-0.008227	0.26	9.6216	0.0030					1560.3	0.5	8.945	7.972	Li computed from Rv vs Log(f)
			6.78	6.0	-1421.223	1617.3	0.97	9.9729	-73.4	100198	60.16	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	100409	-211.2	0.000000	0.00	9.9520	0.0210					1613.9	3.4	8.958	8.090
			6.85	7.0	-1934.442	1694.7	1.02	10.4502	-95.3	95597	57.42	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	96307	-710.6	0.000000	0.00	10.3732	0.0770					1682.2	12.5	8.997	8.311
			6.90	8.0	-2526.619	1789.4	1.07	11.0342	-117.9	90510	54.38	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	91574	-1064.2	0.000000	0.00	10.9061	0.1281					1768.6	20.8	9.023	8.368
			6.95	9.0	-3197.752	1910.0	1.15	11.7778	-139.8	84765	50.94	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	86209	-1444.0	0.000000	0.00	11.5809	0.1970					1878.1	31.9	9.053	8.401
			7.00	10.0	-3947.842	2064.0	1.24	12.7275	-159.7	78411	47.14	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	80214	-1803.5	0.000000	0.00	12.4419	0.2856					2017.7	46.3	9.081	8.409
			7.04	11.0	-4776.889	2263.0	1.36	13.9546	-176.3	71485	43.00	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	73587	-2102.6	0.000000	0.00	13.5568	0.3978					2198.5	64.5	9.104	8.396
			7.08	12.0	-5684.892	2524.0	1.51	15.5640	-188.1	64063	38.55	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	66330	-2267.0	0.000000	0.00	15.0336	0.5304					2438.0	86.0	9.117	8.359
			7.11	13.0	-6671.853	2880.0	1.73	17.7592	-193.5	56115	33.79	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	58441	-2325.7	0.000000	0.00	17.0548	0.7044					2765.8	114.2	9.121	8.313
			7.15	14.0	-7737.770	3397.0	2.04	20.9473	-190.2	47549	28.64	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	49921	-2372.3	0.000000	0.00	19.9556	0.9917					3236.2	160.8	9.125	8.275
			7.18	15.0	-8882.644	4220.0	2.53	26.0222	-175.8	38253	23.06	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	40770	-2517.1	0.000000	0.00	24.4225	1.5997					3960.6	259.4	9.138	8.257
			7.20	16.0	-10106.475	5554.0	3.33	34.2482	-152.0	29047	17.52	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	30988	-1941.2	0.000000	0.00	32.1132	2.1350					5207.8	346.2	9.093	8.170
			7.23	17.0	-11409.263	8574.0	5.14	52.8707	-111.1	18803	11.35	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	20575	-1771.7	0.000000	0.00	48.3424	4.5283					7839.6	734.4	9.082	8.139
			7.26	18.0	-12791.007	20530.0	12.32	126.5963	-52.0	7847	4.74	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	9530	-1683.2	0.000000	0.00	104.3586	22.2377					16923.7	3606.3	9.080	8.121
Derived parameters		esd +/-	sums >>>>>>>>>>>>					16	3.28E-03	-1.181704468	671.135132	357.322830	-126713.778160682			-6.54E-15	15.03
D	8.06E-01
a	111769.22	29.9		y obs - y calc
Li / uH	8.947007	0.0024
b = Ci / pF	7.993030	0.0661
chi^2/(n-2)	1.073870
esd of fit	1.04
Rv0	1450.93
				Incremental error analysis
X-over /MHz	18.82					R0	49.83	49.89	49.95	err contrib
				Li	8.959693	8.947007	8.981270	0.010789
				Ci	6.297673	7.993030	6.282543	-0.007565

				C2	10.1	10.3	10.5
				Li	8.796297	8.947007	9.144665	0.174184
				Ci	6.414655	7.993030	6.170287	-0.122184

				Accuracy
				esd Li	0.174534
				esd Ci	0.139136

				Li' vs Log(f)

## Sheet2
Variation of C1'm (nominal capacitance, measured @ 1.5915MHz) with frequency					To see details of fitting calculation show columns N-P
			Log(f)	f	Dial	Ci'm	esd C1'm	y	esd y	fit	weight	x	wx	wy	wxy	wx^2	y calc	e	we	we^2		<Ci'm>	obs-calc
				/MHz		/pF		1/C1'm			w					obs-calc
see sheet 4 for capcal			6.15	1.40	6.56	109.4	0.08	0.009145	0.000007	1	2.23E+10	-22.62	-505438640876.2	204316644.48	-4622112731.94	1.14E+13	0.009145	0.000000	-10627.4	0.005		109.35	0.006
			6.18	1.50	6.56	109.4	0.08	0.009145	0.000007	1	2.23E+10	-11.17	-249645208667.3	204316644.48	-2282944365.81	2.79E+12	0.009146	-0.000001	-26388.7	0.031		109.34	0.014
			6.20	1.60	6.56	109.4	0.08	0.009145	0.000007	1	2.23E+10	1.06	23789149900.8	204316644.48	217545956.61	2.53E+10	0.009147	-0.000002	-43237.0	0.084		109.33	0.023
			6.23	1.70	6.56	109.4	0.08	0.009145	0.000007	1	2.23E+10	14.09	314864434828.2	204316644.48	2879358235.32	4.44E+12	0.009147	-0.000003	-61172.2	0.167		109.32	0.033
			6.26	1.80	6.56	109.4	0.08	0.009145	0.000007	1	2.23E+10	27.91	623580646114.8	204316644.48	5702492470.31	1.74E+13	0.009148	-0.000004	-80194.4	0.288		109.31	0.043
			6.28	1.90	6.56	109.4	0.08	0.009145	0.000007	1	2.23E+10	42.52	949937783760.6	204316644.48	8686948661.59	4.04E+13	0.009149	-0.000004	-100303.6	0.450		109.30	0.054
			6.30	2.00	6.56	109.4	0.08	0.009145	0.000007	1	2.23E+10	57.91	1293935847765.7	204316644.48	11832726809.15	7.49E+13	0.009150	-0.000005	-121499.8	0.661		109.29	0.065
			6.34	2.20	6.56	109.4	0.08	0.009145	0.000007	1	2.23E+10	91.08	2034854754853.5	204316644.48	18608248973.13	1.85E+14	0.009152	-0.000007	-167153.2	1.251		109.26	0.089
			6.38	2.40	6.56	109.4	0.08	0.009145	0.000007	1	2.23E+10	127.4	2846337367378.3	204316644.48	26029058962.26	3.63E+14	0.009154	-0.000010	-217154.4	2.111		109.24	0.116
			6.41	2.60	6.55	109.1	0.08	0.009163	0.000007	1	2.22E+10	166.87	3699422292254.0	203125164.59	33896329639.39	6.17E+14	0.009157	0.000006	126219.3	0.719		109.21	-0.068
			6.45	2.80	6.55	109.1	0.08	0.009163	0.000007	1	2.22E+10	209.51	4644632617640.4	203125164.59	42556914519.07	9.73E+14	0.009160	0.000003	67978.1	0.208		109.18	-0.037
			6.48	3.00	6.55	109.1	0.08	0.009163	0.000007	1	2.22E+10	255.31	5659858522685.1	203125164.59	51859024204.65	1.44E+15	0.009162	0.000000	5422.7	0.001		109.14	-0.003
			6.54	3.50	6.54	108.9	0.08	0.009181	0.000007	1	2.20E+10	383.61	8438080784022.3	201938814.8	77465873068.05	3.24E+15	0.009170	0.000010	225389.3	2.309		109.05	-0.122
			6.60	4.00	6.54	108.9	0.08	0.009181	0.000007	1	2.20E+10	531.65	11694528175378.5	201938814.8	107361716296.91	6.22E+15	0.009179	0.000001	24736.2	0.028		108.94	-0.013
			6.65	4.50	6.53	108.7	0.08	0.009198	0.000007	1	2.18E+10	699.44	15265292458524.8	200757581.86	140417472806.39	1.07E+16	0.009190	0.000009	191115.3	1.674		108.82	-0.104
			6.70	5.00	6.53	108.7	0.08	0.009198	0.000007	1	2.18E+10	886.96	19357986497309.1	200757581.86	178064033162.65	1.72E+16	0.009201	-0.000003	-61065.0	0.171		108.68	0.033
			6.78	6.00	6.51	108.3	0.08	0.009235	0.000007	1	2.15E+10	1321.22	28387168652936.2	198410413.71	262144408734.48	3.75E+16	0.009228	0.000007	141249.9	0.929		108.37	-0.077
			6.85	7.00	6.50	108.1	0.08	0.009253	0.000007	1	2.13E+10	1834.44	39105437416940.5	197244452.08	361833598406.46	7.17E+16	0.009260	-0.000007	-146847.7	1.012		108.00	0.080
			6.90	8.00	6.48	107.7	0.08	0.009289	0.000007	1	2.10E+10	2426.62	50920536440049.5	194927707.67	473015225791.19	1.24E+17	0.009296	-0.000007	-144041.9	0.989		107.57	0.079
			6.95	9.00	6.45	107.0	0.08	0.009345	0.000007	1	2.05E+10	3097.75	63479832833293.7	191490340.36	593189551517.5	1.97E+17	0.009337	0.000007	144179.0	1.014		107.10	-0.081
			7.00	10.00	6.42	106.4	0.08	0.009400	0.000007	1	2.00E+10	3847.84	76993836964912.6	188097955.92	723771169845.26	2.96E+17	0.009384	0.000017	333357.0	5.554		106.57	-0.189
			7.04	11.00	6.40	106.0	0.08	0.009438	0.000007	1	1.97E+10	4676.89	92101902432462.5	185861181.25	869252026769.03	4.31E+17	0.009435	0.000003	61759.8	0.194		105.99	-0.035
			7.08	12.00	6.37	105.3	0.08	0.009495	0.000007	1	1.92E+10	5584.89	107372968075053.0	182542978.7	1019482846042.94	6.00E+17	0.009491	0.000004	77497.9	0.312		105.37	-0.045
			7.11	13.00	6.34	104.7	0.08	0.009552	0.000007	1	1.88E+10	6571.85	123335318254444.0	179268812.34	1178128206005.18	8.11E+17	0.009552	0.000001	12632.3	0.009		104.69	-0.007
			7.15	14.00	6.31	104.1	0.08	0.009610	0.000007	1	1.83E+10	7637.77	139905956749950.0	176038328.05	1344540234466.99	1.07E+18	0.009617	-0.000007	-126912.3	0.879		103.98	0.075
			7.18	15.00	6.27	103.2	0.08	0.009689	0.000008	1	1.77E+10	8782.64	155731787352415.0	171798354.42	1508843779970.08	1.37E+18	0.009688	0.000001	16786.8	0.016		103.22	-0.010
			7.20	16.00	6.23	102.4	0.08	0.009768	0.000008	1	1.72E+10	10006.47	171721981057961.0	167634572.43	1677431142511.13	1.72E+18	0.009763	0.000005	87525.5	0.446		102.43	-0.053
			7.23	17.00	6.20	101.7	0.08	0.009829	0.000008	1	1.67E+10	11309.26	189349573472395.0	164561234.25	1861066226312.59	2.14E+18	0.009843	-0.000015	-246753.5	3.637		101.59	0.152
			7.26	18.00	6.15	100.7	0.08	0.009931	0.000008	1	1.61E+10	12691.01	203870342605680.0	159532244.53	2024624880580.47	2.59E+18	0.009929	0.000002	37502.2	0.088		100.72	-0.024
			sums >>>>>>>					29	6.03E+11		1518368659791370.0	5611026663.12	14595995983621	1.15E+19			-8.70E-07	25.235
Derived parameters		esd					Chi^2(/n-2) >>>>>		0.935
Denom	4.63E+30					esd of fit >>>>>>		0.967
a	0.009147	0.000002
C1' /pF	109.330	0.018					Ci'measured vs Log(f) (with expectation value)
b	6.16E-08	3.49E-10
L1 /nH	61.617	0.349

KT	11.63	< Xformer const. = ( N R0 / k Ri) + (1/N) -1
L2 /nH	50	20
L1(true)	65.917	1.755	nH	< corrected for L2

				C1'm obs-calc vs Log(f)

## Sheet3
Stray capacitance of voltage sampling network				Predicted C1 = C1' + C1s = C2 [( N R0 / k Ri) + (1/N) -1]
Parameter		esd +/-	Deriv dC1/dpar	(der*esd)^2
N =	12
R0 =	49.89	0.06	2.59	0.02
C2 /pF =	10.3	0.2	11.63	5.41
Ri =	49.76	0.06	-2.6	0.02
k =	0.959	0.0112	-134.75	2.28
				7.74	<<< sum, used to calculate esd of predicted C1.
KT	11.629	< transformer const = ( N R0 / k Ri) + (1/N) -1

C1' =	109.330	1.093
C1 pred =	119.779	2.781
C1s =	10.449	2.988

C1 Prediction for bridge evaluation (see sheet 5)
R0 =	50
C1 pred =	120.06

Self capacitance of transformer secondary				Ci' = Ci apparent + ( Csh * Ro / Ri ) + [ L2 / ( N Ro Rv ) ] + ( LH / Ri^2 ) - Cis - [ L0 / ( R0 Ri ) ] + C0 Ro / Ri
parameter		esd +/-
Csh / pF	0.00	0.00	< shield protrusion capacitance. Set to 0 for gen. side earth.
L2 / nH	50.00	20.00	< inductance of upper voltage sampling arm
L2 fit	1		< fitting flag =1 when L2 is included in fit on sheet 1, 0 otherwise.
RV0	1450.93		< optimal Rv
LH / nH	6.00	2.00	< inductance of sec. load resistor
Cis / pF	0.40	0.10	< strays accross sec winding
L0 / nH	13.00	3.00	< inductance of pri. transmission line
C0 / pF	3.50	0.50	< capacitance of pri. transmission line

Series for Ci'		esd +/-	Variance
Ci apparent	7.99	0.07	0.004	< from fit on sheet 1
Corr for Csh	0.00	0.00	0.000
Corr for L2	0.00	0.02	0.001	< inclusion depends on whether L2 is included in fit on sheet 1.
Corr for LH	2.42	0.81	0.652
Corr for Cis	-0.40	0.10	0.010
Corr for L0	-5.24	1.21	1.460
Corr for C0	3.51	0.50	0.251
Ci' / pF	8.29	1.54	2.38	Only valid if parameters have been filled in correctly!

lw / mm	228.00	< length of secondary winding wire
Ci' pred / pF	7.64	< predicted Ci' if phase velocity = c

## Sheet4
Data reduction for 500pF ref. cap.					C =	69.83	*exp( -	0.32	* x) -	56.91325	+	24.040788	* x
Non linearity at low settings is corrected by subtracting an exponential function f(x) =( C0-a)Exp(-px)					where x is the dial reading
seed for C0	12.92	<<< seed values are required because the fitting process is iterative
seed for -a	56.91	<<<
C0 - a	69.83					Show columns J-M for details of fitting calculation.
p	0.32
		Dial	f(x)	C meas	esd C	y	fit	w	wy	wx	wx^2	wxy	y calc	obs-calc	we	we^2		C calc	obs-calc	w ec	w (ec)^2
		x		/ pF		C - f(x)					e					ec
0 - 25pF range		0	69.83	15.05	1.06	-54.8	1	0.890	-48.74	0.00	0.00	0.00	-56.91	2.13	1.90	4.049		12.92	2.13	1.898	4.049
		1	50.71	17.40	0.87	-33.3	1	1.322	-44.04	1.32	1.32	-44.04	-32.87	-0.43	-0.57	0.250		17.83	-0.43	-0.575	0.250
0 - 250pF range		2	36.82	28.50	0.91	-8.3	1	1.197	-9.96	2.39	4.79	-19.92	-8.83	0.51	0.61	0.312		27.99	0.51	0.611	0.312
		3	26.74	41.50	0.97	14.8	1	1.072	15.82	3.21	9.64	47.46	15.21	-0.45	-0.48	0.214		41.95	-0.45	-0.479	0.214
		4	19.42	58.00	1.03	38.6	1	0.939	36.23	3.76	15.02	144.92	39.25	-0.67	-0.62	0.416		58.67	-0.67	-0.625	0.416
		5	14.10	76.00	1.10	61.9	1	0.820	50.79	4.10	20.51	253.94	63.29	-1.39	-1.14	1.583		77.39	-1.39	-1.140	1.583
		6	10.24	96.00	1.18	85.8	1	0.713	61.18	4.28	25.68	367.07	87.33	-1.57	-1.12	1.756		97.57	-1.57	-1.119	1.756
		7	7.43	118.00	1.27	110.6	1	0.618	68.34	4.33	30.28	478.35	111.37	-0.81	-0.50	0.402		118.81	-0.81	-0.498	0.402
		8	5.40	141.00	1.36	135.6	1	0.537	72.88	4.30	34.40	583.08	135.41	0.19	0.10	0.019		140.81	0.19	0.101	0.019
		9	3.92	163.00	1.45	159.1	1	0.474	75.45	4.27	38.42	679.09	159.45	-0.37	-0.18	0.066		163.37	-0.37	-0.177	0.066
		10	2.85	187.00	1.55	184.2	1	0.417	76.85	4.17	41.73	768.49	183.49	0.66	0.27	0.181		186.34	0.66	0.275	0.181
		11	2.07	212.00	1.65	209.9	1	0.368	77.30	4.05	44.55	850.27	207.54	2.40	0.88	2.117		209.60	2.40	0.883	2.117
		12	1.50	236.50	1.75	235.0	1	0.328	77.09	3.94	47.24	925.04	231.58	3.42	1.12	3.843		233.08	3.42	1.123	3.843
0 - 2500pF range		13	1.09	260.00	3.54	258.9	1	0.080	20.66	1.04	13.49	268.59	255.62	3.29	0.26	0.865		256.71	3.29	0.263	0.865
		14	0.79	280.00	3.62	279.2	1	0.076	21.31	1.07	14.96	298.29	279.66	-0.45	-0.03	0.015		280.45	-0.45	-0.034	0.015
		15	0.57	302.00	3.71	301.4	1	0.073	21.92	1.09	16.36	328.85	303.70	-2.27	-0.17	0.376		304.27	-2.27	-0.165	0.376
		16	0.42	327.00	3.81	326.6	1	0.069	22.52	1.10	17.65	360.35	327.74	-1.16	-0.08	0.092		328.16	-1.16	-0.080	0.092
		17	0.30	350.00	3.90	349.7	1	0.066	22.99	1.12	19.00	390.85	351.78	-2.08	-0.14	0.285		352.08	-2.08	-0.137	0.285
		18	0.22	375.00	4.00	374.8	1	0.063	23.42	1.13	20.25	421.63	375.82	-1.04	-0.07	0.068		376.04	-1.04	-0.065	0.068
		19	0.16	400.00	4.10	399.8	1	0.059	23.79	1.13	21.48	451.93	399.86	-0.02	0.00	0.000		400.02	-0.02	-0.001	0.000
		20	0.12	425.00	4.20	424.9	1	0.057	24.09	1.13	22.68	481.73	423.90	0.98	0.06	0.055		424.02	0.98	0.056	0.055
		21	0.08	448.00	4.29	447.9	1	0.054	24.32	1.14	23.94	510.62	447.94	-0.03	0.00	0.000		448.03	-0.03	-0.001	0.000
		22	0.06	470.00	4.38	469.9	1	0.052	24.50	1.15	25.23	538.91	471.98	-2.05	-0.11	0.218		472.05	-2.05	-0.107	0.218
		23	0.04	498.00	4.49	498.0	1	0.050	24.68	1.14	26.22	567.60	496.02	1.93	0.10	0.185		496.07	1.93	0.096	0.185
		24	0.03	520.00	4.58	520.0	1	0.048	24.79	1.14	27.46	594.92	520.07	-0.10	0.00	0.000		520.10	-0.10	-0.005	0.000
		25	0.02	542.00	4.67	542.0	1	0.046	24.87	1.15	28.68	621.81	544.11	-2.13	-0.10	0.208		544.13	-2.13	-0.098	0.208
		sums >>>>>>>>>					26	10.489	813.03	58.65	590.98	10869.79			0.00	17.576				0.000	17.576
D	2758.883
a	-56.913247
esd a	0.156871
b	24.040788
esd b	0.002784
Chi^2/n-2)	0.732319
esd fit	0.855756

## Sheet5
Bridge performance evaluation - change Rvset and Cvset to see the effect of different calibration choices.
N	12
R0	50
C2 /pF	10.3
Li /uH	8.947
Rv0	1447.74	<<< Guide value (estimated value for zero phase error at 0Hz).
Rvset	1455.31	<<< selected value (usually set to balance bridge at some chosen 'calibration' frequency, but can be arbitrary)
C1s /pF	10.449
L1 /nH	61.617
C1 pred	120.064					Show columns W-Z for details of fit.
Cv set 0	130.364	<<< Guide value (estimated value for zero magnitude error to zero at 0Hz)
C1 set	120.000
Cv set	130.300	<<< selected value (see note for Rvset)
			x					y				Determining the gradient of Phi vs. f.
			f	Rv	C1'm	C1'0	Cv	Xcv	Xcvset	U	Rbal	Xbal	|Zbal|	(|Zbal|/R0)-1	Phi				esd Phi	fit	weight					Phi calc	obs-calc
		Log(f)	/MHz	Ohms	/pF	/pF	/pF					Magn. err.	deg.				esd y		w	wx	wy	wxy	wx^2	y calc	e	we	we^2
		6.15	1.4	1462.6	109.35	109.28	130.03	-874.29	-872.46	3.47E-07	50.14	-0.06	50.14	0.003	-0.073				0.0500	1	400	560	-29	-41	784	-0.038	-0.035	-14.2	0.503
		6.18	1.5	1462.6	109.35	109.28	130.03	-816.00	-814.30	3.59E-07	50.14	-0.06	50.14	0.003	-0.071				0.0200	1	2500	3750	-177	-266	5625	-0.053	-0.018	-45.9	0.844
		6.20	1.6	1462.5	109.35	109.28	130.03	-765.00	-763.41	3.70E-07	50.14	-0.06	50.14	0.003	-0.067				0.0075	1	17778	28444	-1189	-1902	45511	-0.067	0.000	5.5	0.002
		6.23	1.7	1464.2	109.35	109.28	130.03	-720.00	-718.50	3.79E-07	50.14	-0.08	50.14	0.003	-0.091				0.0075	1	17778	30222	-1616	-2747	51378	-0.082	-0.009	-161.7	1.471
		6.26	1.8	1464.7	109.35	109.28	130.03	-680.00	-678.58	3.87E-07	50.14	-0.08	50.14	0.003	-0.095				0.0075	1	17778	32000	-1692	-3045	57600	-0.096	0.001	22.1	0.028
Calibration freq >>		6.28	1.9	1466.9	109.35	109.28	130.03	-644.21	-642.87	3.94E-07	50.15	-0.11	50.15	0.003	-0.123				0.0075	1	17778	33778	-2195	-4171	64178	-0.111	-0.012	-221.9	2.771
		6.30	2.0	1467.1	109.35	109.28	130.03	-612.00	-610.73	4.00E-07	50.15	-0.11	50.15	0.003	-0.122				0.0075	1	17778	35556	-2169	-4337	71111	-0.126	0.004	64.7	0.236
		6.34	2.2	1469.5	109.35	109.28	130.03	-556.37	-555.20	4.10E-07	50.15	-0.13	50.15	0.003	-0.145				0.0075	1	17778	39111	-2575	-5666	86044	-0.155	0.010	177.5	1.772
		6.38	2.4	1473.0	109.35	109.28	130.03	-510.00	-508.94	4.17E-07	50.16	-0.16	50.16	0.003	-0.178				0.0075	1	17778	42667	-3159	-7582	102400	-0.184	0.006	113.2	0.721
		6.41	2.6	1478.2	109.14	109.07	129.82	-471.54	-469.79	4.23E-07	50.24	-0.17	50.24	0.005	-0.198				0.0075	1	17778	46222	-3521	-9153	120178	-0.213	0.015	271.6	4.150
		6.45	2.8	1483.2	109.14	109.07	129.82	-437.86	-436.23	4.28E-07	50.25	-0.21	50.25	0.005	-0.239				0.0075	1	17778	49778	-4247	-11891	139378	-0.243	0.004	64.7	0.236
		6.48	3.0	1488.7	109.14	109.07	129.82	-408.67	-407.15	4.31E-07	50.25	-0.25	50.25	0.005	-0.280				0.0075	1	17778	53333	-4970	-14909	160000	-0.272	-0.008	-138.5	1.078
		6.54	3.5	1501.9	108.93	108.85	129.60	-350.86	-348.99	4.36E-07	50.34	-0.29	50.34	0.007	-0.336				0.0075	1	17778	62222	-5969	-20892	217778	-0.345	0.009	161.1	1.459
		6.60	4.0	1518.7	108.93	108.85	129.60	-307.01	-305.36	4.37E-07	50.34	-0.37	50.35	0.007	-0.422				0.0075	1	17778	71111	-7501	-30003	284444	-0.418	-0.004	-71.4	0.287
		6.65	4.5	1538.9	108.71	108.64	129.39	-273.34	-271.43	4.36E-07	50.43	-0.43	50.43	0.009	-0.493				0.0075	1	17778	80000	-8765	-39445	360000	-0.491	-0.002	-37.1	0.078
		6.70	5.0	1560.8	108.71	108.64	129.39	-246.01	-244.29	4.33E-07	50.43	-0.50	50.43	0.009	-0.572				0.0075	1	17778	88889	-10166	-50829	444444	-0.564	-0.008	-138.5	1.080
		6.78	6.0	1617.3	108.29	108.22	128.97	-205.68	-203.58	4.22E-07	50.60	-0.63	50.60	0.012	-0.715				0.0075	1	17778	106667	-12718	-76306	640000	-0.710	-0.005	-92.3	0.480
		6.85	7.0	1694.7	108.08	108.00	128.75	-176.59	-174.49	4.06E-07	50.68	-0.79	50.69	0.014	-0.888				0.0075	0	0	0	0	0	0	-0.856	-0.032	0.0	0.000
		6.90	8.0	1789.4	107.65	107.58	128.33	-155.03	-152.68	3.87E-07	50.85	-0.92	50.86	0.017	-1.038				0.0075	0	0	0	0	0	0	-1.002	-0.035	0.0	0.000
		6.95	9.0	1910.0	107.01	106.94	127.69	-138.49	-135.72	3.65E-07	51.10	-1.05	51.11	0.022	-1.181				0.0075	0	0	0	0	0	0	-1.149	-0.032	0.0	0.000
		7.00	10.0	2064.0	106.38	106.31	127.06	-125.26	-122.15	3.40E-07	51.35	-1.19	51.36	0.027	-1.325				0.0075	0	0	0	0	0	0	-1.295	-0.030	0.0	0.000
		7.04	11.0	2263.0	105.96	105.89	126.64	-114.25	-111.04	3.12E-07	51.51	-1.32	51.53	0.031	-1.473				0.0075	0	0	0	0	0	0	-1.441	-0.032	0.0	0.000
		7.08	12.0	2524.0	105.32	105.25	126.00	-105.26	-101.79	2.81E-07	51.77	-1.46	51.79	0.036	-1.613				0.0075	0	0	0	0	0	0	-1.587	-0.026	0.0	0.000
		7.11	13.0	2880.0	104.69	104.62	125.37	-97.65	-93.96	2.48E-07	52.02	-1.59	52.04	0.041	-1.752				0.0075	0	0	0	0	0	0	-1.733	-0.019	0.0	0.000
		7.15	14.0	3397.0	104.05	103.99	124.74	-91.14	-87.25	2.11E-07	52.28	-1.73	52.30	0.046	-1.894				0.0075	0	0	0	0	0	0	-1.879	-0.015	0.0	0.000
		7.18	15.0	4220.0	103.21	103.15	123.90	-85.64	-81.43	1.71E-07	52.62	-1.87	52.66	0.053	-2.040				0.0075	0	0	0	0	0	0	-2.025	-0.015	0.0	0.000
		7.20	16.0	5554.0	102.37	102.31	123.06	-80.83	-76.34	1.31E-07	52.97	-2.01	53.01	0.060	-2.169				0.0075	1	17778	284444	-38559	-616947	4551111	-2.172	0.003	46.3	0.121
		7.23	17.0	8574.0	101.74	101.68	122.43	-76.47	-71.85	8.53E-08	53.23	-2.15	53.28	0.066	-2.315				0.0075	1	17778	302222	-41164	-699782	5137778	-2.318	0.002	39.9	0.089
		7.26	18.0	20530.0	100.70	100.63	121.38	-72.84	-67.86	3.59E-08	53.68	-2.31	53.73	0.075	-2.466				0.0075	1	17778	320000	-43847	-789237	5760000	-2.464	-0.003	-45.0	0.114
					Sums >>>>		20	322900	1710976.67	-196226.8	-2389152.52	18299742.33			0	17.52
	Phase error vs f					Parameter		+/- esd					Chi^2/(n-2) >>>>		0.973159
					D	2.98E+12					esd of an obs of unit w >>>>			0.986
					a	0.166653	0.002444	phase intercept at 0Hz
					b	-0.146138	0.000325	degrees / MHz
					tp	0.405940	0.000902	ns (propagation delay)

					Phase err. Deviation from linearity vs. Log(f)

		Magnitude error vs f