## Sheet1
FT50-61, 12 turns 0.9mm wire + 1 turn shift comp with 3.34 Ohm load. Lsec = 9.25uH @ 1.5915MHz. 40pF variable ref cap with padding. No shield. L-bal coil.					D. W. Knight. June 2007.			Version	1.01
		esd +/-					Show columns N-Q to see deatails of fitting procedure.					See sheet 2 for C1'm vs frequency
N =	12.00					Li' = N R0 C2 Rv
R0 =	49.89	0.06					from fit	from fit
C2+Cx/pF	10.90	0.20					Ci =	Li =
L2 / nH	900.00	20.00					Li computed from Rv
L2 fit	1	< 1 to include L2 in fit, 0 to exclude it.
1/(N R0 C2)	1.53E+08					6.046	9.627
			Log (f)	f	x	Rv	esd Rv	Li'	Corr(L2)	y	esd y	fit	weight	w x	w x^2	w y	w x y	y calc	obs-calc	w e	w e^2	Li' calc	Li'					Computed	Computed
				/MHz	-(2 pi f)^2	Ohms	Ohms	/ uH					w					e			/ uH	obs-calc					Li /uH	Ci /pF
			6.20	1.59	-100.000	1483.2	14.83	9.6788	-101.4	103217.37	1033.2	1	9.37E-07	-0.000093679	0.009368	0.096693	-9.669299736	103269.28	-51.92	-0.000049	0.00	9.6739	0.0049					9.622	5.552
			6.23	1.70	-114.093	1482.6	7.41	9.6749	-115.7	103244.85	516.8	1	3.74E-06	-0.000427177	0.048738	0.386562	-44.103845568	103184.08	60.77	0.000228	0.01	9.6806	-0.0057					9.611	4.499
			6.26	1.80	-127.910	1483.2	0.89	9.6788	-129.6	103189.08	62.0	1	2.60E-04	-0.033284664	4.257444	26.851786	-3434.613850444	103100.54	88.54	0.023040	2.04	9.6871	-0.0083					9.607	4.340
			6.28	1.90	-142.517	1485.2	0.89	9.6918	-144.3	103035.34	61.9	1	2.61E-04	-0.037185773	5.299608	26.884136	-3831.448722028	103012.22	23.12	0.006032	0.14	9.6940	-0.0022					9.612	4.872
Calibration freq. >>>			6.30	2.00	-157.914	1485.8	0.89	9.6958	-159.8	102978.15	61.9	1	2.61E-04	-0.041236370	6.511787	26.890929	-4246.445235985	102919.13	59.02	0.015412	0.91	9.7013	-0.0056					9.607	4.661
			6.34	2.20	-191.076	1488.8	0.89	9.7153	-192.9	102737.16	61.8	1	2.62E-04	-0.050097703	9.572446	26.936445	-5146.895823948	102718.63	18.53	0.004858	0.09	9.7171	-0.0017					9.607	4.939
			6.38	2.40	-227.396	1491.7	0.90	9.7343	-229.2	102500.83	61.6	1	2.63E-04	-0.059852900	13.610291	26.979280	-6134.971812797	102499.04	1.79	0.000472	0.00	9.7344	-0.0002					9.606	5.030
			6.41	2.60	-266.874	1494.4	0.90	9.7519	-268.5	102275.92	61.5	1	2.64E-04	-0.070498544	18.814236	27.017622	-7210.303565673	102260.35	15.58	0.004115	0.06	9.7534	-0.0015					9.601	4.982
			6.45	2.80	-309.511	1498.6	0.90	9.7793	-310.5	101946.51	61.4	1	2.66E-04	-0.082221851	25.448550	27.082193	-8382.230904541	102002.56	-56.05	-0.014889	0.83	9.7739	0.0054					9.604	5.224
			6.48	3.00	-355.306	1501.9	0.90	9.8008	-355.6	101676.67	61.2	1	2.67E-04	-0.094803479	33.684222	27.129598	-9639.302542845	101725.68	-49.00	-0.013075	0.64	9.7961	0.0047					9.599	5.183
			6.54	3.50	-483.611	1512.5	0.91	9.8700	-480.7	100836.57	60.8	1	2.71E-04	-0.130865927	63.288152	27.286563	-13196.071630153	100949.93	-113.36	-0.030676	3.48	9.8590	0.0110					9.593	5.287
			6.60	4.00	-631.655	1523.6	0.91	9.9424	-623.2	99955.87	60.3	1	2.75E-04	-0.173444943	109.557310	27.446705	-17336.839885338	100054.84	-98.97	-0.027177	2.69	9.9326	0.0098					9.579	5.216
			6.65	4.50	-799.438	1535.6	0.92	10.0207	-782.6	99010.50	59.9	1	2.79E-04	-0.222987730	178.264855	27.617062	-22078.127677431	99040.41	-29.90	-0.008341	0.25	10.0177	0.0030					9.558	5.105
			6.70	5.00	-986.960	1547.8	0.93	10.1003	-958.6	98047.96	59.4	1	2.83E-04	-0.279685161	276.038190	27.784861	-27422.558510719	97906.63	141.33	0.040051	5.66	10.1148	-0.0144					9.526	4.932
			6.78	6.00	-1421.223	1577.3	0.95	10.2928	-1354.6	95800.29	58.3	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	95281.03	519.26	0.000000	0.00	10.3482	-0.0553					9.456	4.728
			6.85	7.00	-1934.442	1617.6	0.97	10.5558	-1797.8	92936.62	56.8	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	92178.04	758.57	0.000000	0.00	10.6410	-0.0852					9.396	4.725
			6.90	8.00	-2526.619	1657.6	0.99	10.8169	-2291.4	90156.88	55.5	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	88597.68	1559.20	0.000000	0.00	11.0024	-0.1856					9.283	4.522
			6.95	9.00	-3197.752	1693.5	1.02	11.0511	-2838.6	87649.91	54.3	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	84539.94	3109.98	0.000000	0.00	11.4445	-0.3933					9.106	4.186	Ci computed from Rv
			7.00	10.00	-3947.842	1733.4	1.04	11.3115	-3423.8	84981.83	53.0	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	80004.81	4977.02	0.000000	0.00	11.9863	-0.6748					8.907	3.918
			7.08	12.00	-5684.892	1801.3	1.08	11.7546	-4744.4	80328.76	51.0	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	69502.41	10826.34	0.000000	0.00	13.4686	-1.7140					8.372	3.307
			7.15	14.00	-7737.770	1859.0	1.12	12.1311	-6257.3	76175.41	49.5	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	57090.49	19084.92	0.000000	0.00	15.7859	-3.6548					7.739	2.771
			7.20	16.00	-10106.475	1880.2	1.13	12.2695	-8080.6	73422.61	48.9	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	42769.04	30653.57	0.000000	0.00	19.6658	-7.3964					7.012	2.214
			7.26	18.00	-12791.007	1859.4	1.12	12.1337	-10341.4	72073.53	49.4	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	26538.06	45535.47	0.000000	0.00	27.1154	-14.9816					6.260	1.678
			7.30	20.00	-15791.367	1828.7	1.10	11.9334	-12981.5	70817.00	50.3	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	8397.56	62419.45	0.000000	0.00	46.7747	-34.8414					5.578	1.271
			7.34	22.00	-19107.554	1762.5	1.06	11.5014	-16297.6	70648.40	52.2	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	-11652.47	82300.87	0.000000	0.00	215.2793	-203.7779					4.939	0.886
			7.38	24.00	-22739.569	1680.8	1.01	10.9682	-20338.3	70833.99	54.7	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	-33612.03	104446.02	0.000000	0.00	-75.3366	86.3048					4.373	0.559
			7.41	26.00	-26687.410	1583.2	0.95	10.3313	-25340.7	71452.09	58.1	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	-57481.12	128933.21	0.000000	0.00	-31.1135	41.4448					3.874	0.265
Calibration freq. >>>			7.45	28.00	-30951.079	1476.3	0.89	9.6338	-31517.3	72284.33	62.3	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	-83259.73	155544.05	0.000000	0.00	-19.3265	28.9603					3.437	0.002
			7.48	30.00	-35530.576	1347.2	0.81	8.7913	-39647.7	74101.07	68.2	0	0.00E+00	0.000000000	0.000000	0.000000	0.000000000	-110947.87	185048.93	0.000000	0.00	-14.0252	22.8165					3.043	-0.278
Derived parameters		esd +/-	sums >>>>>>>>>>>>					14	3.22E-03	-1.276685902	744.405196	326.390433	-128113.583307206			-8.22E-14	16.81
D	7.64E-01
a	103873.90	36.94
Li / uH	9.627058	0.0034		y obs - y calc
b = Ci / pF	6.046110	0.0768
chi^2/(n-2)	1.401
esd of fit	1.18
Rv0	1475.27
				Incremental error analysis
X-over /MHz	20.86					R0	49.83	49.89	49.95	err contrib.
				Li	8.994662	9.627058	9.016323	0.010831
				Ci	6.173432	6.046110	6.158601	-0.007416

				C2	10.1	10.3	10.5
				Li	8.828914	9.627058	9.182070	0.176578
				Ci	6.289328	6.046110	6.047431	-0.120949

				Accuracy
				esd Li	0.176943
				esd Ci	0.143453

				Li' vs Log(f)

## Sheet2
Variation of C1'm (nominal capacitance, measured @ 1.5915MHz) with frequency					To see details of fitting calculation show columns N-P
			Log(f)	f	Dial	Ci'm	esd C1'm	y	esd y	fit	weight	x	wx	wy	wxy	wx^2	y calc	e	we	we^2		<Ci'm>	obs-calc
				/MHz		/pF		1/C1'm			w					obs-calc
			6.20	1.59	12.050	117.55	0.63	0.008507	0.000045	1	4.89E+08	0.0	34951	4157821	297.34	2.50E+00	0.008480	0.000028	13488.9	0.372		117.93	-0.383
			6.23	1.70	12.055	117.56	0.13	0.008506	0.000009	1	1.22E+10	14.1	172263838400	103979235	1465340563	2.43E+12	0.008480	0.000027	326253.3	8.708		117.93	-0.370
			6.26	1.80	12.050	117.55	0.10	0.008507	0.000007	1	1.91E+10	27.9	532838213154	162414871	4533010924	1.49E+13	0.008480	0.000028	527299.6	14.564		117.93	-0.383
			6.28	1.90	12.055	117.56	0.10	0.008506	0.000007	1	1.91E+10	42.5	812055477152	162467555	6907647254	3.45E+13	0.008480	0.000027	510165.6	13.627		117.93	-0.370
			6.30	2.00	12.050	117.55	0.08	0.008507	0.000006	1	2.98E+10	57.9	1727569414427	253773237	14696939586	1.00E+14	0.008480	0.000028	824556.5	22.792		117.93	-0.383
			6.34	2.20	12.050	117.55	0.08	0.008507	0.000006	1	2.98E+10	91.1	2716790668839	253773237	23112534869	2.47E+14	0.008480	0.000028	825276.0	22.832		117.93	-0.384
			6.38	2.40	12.055	117.56	0.08	0.008506	0.000006	1	2.98E+10	127.4	3801867159940	253855554	32340102355	4.84E+14	0.008480	0.000027	798976.1	21.391		117.93	-0.371
			6.41	2.60	12.070	117.60	0.08	0.008504	0.000006	1	2.99E+10	166.9	4986484205902	254102615	42403145921	8.32E+14	0.008480	0.000024	718464.2	17.274		117.93	-0.333
			6.45	2.80	12.075	117.61	0.08	0.008503	0.000006	1	2.99E+10	209.5	6263248274152	254185004	53254501993	1.31E+15	0.008480	0.000023	692232.9	16.029		117.93	-0.321
			6.48	3.00	12.090	117.65	0.08	0.008500	0.000006	1	2.99E+10	255.3	7642173208534	254432278	64958025686	1.95E+15	0.008480	0.000020	611649.9	12.498		117.93	-0.284
			6.54	3.50	12.110	117.70	0.08	0.008496	0.000006	1	3.00E+10	383.6	11502634534545	254762226	97729494368	4.41E+15	0.008479	0.000017	505428.1	8.519		117.93	-0.234
			6.60	4.00	12.135	117.76	0.08	0.008492	0.000006	1	3.00E+10	531.7	15976216582098	255175062	135665016418	8.49E+15	0.008479	0.000012	371995.5	4.605		117.93	-0.172
			6.65	4.50	12.055	117.56	0.08	0.008506	0.000006	1	2.98E+10	699.4	20873314420469	253855554	177556210264	1.46E+16	0.008479	0.000027	811392.7	22.061		117.94	-0.377
			6.70	5.00	12.075	117.61	0.08	0.008503	0.000006	1	2.99E+10	887.0	26515356747068	254185004	225452042920	2.35E+16	0.008479	0.000024	706962.9	16.719		117.94	-0.328
			6.78	6.00	12.130	117.75	0.08	0.008493	0.000006	1	3.00E+10	1321.2	39685600420340	255092459	337034032887	5.24E+16	0.008479	0.000014	416612.9	5.778		117.94	-0.193
			6.85	7.00	12.190	117.90	0.08	0.008482	0.000006	1	3.02E+10	1834.4	55387187795446	256084871	469772961237	1.02E+17	0.008478	0.000003	98384.7	0.321		117.95	-0.045
			6.90	8.00	12.310	118.21	0.08	0.008460	0.000006	1	3.05E+10	2426.6	74027802357141	258077409	626255474447	1.80E+17	0.008478	-0.000018	-555072.6	10.100		117.95	0.254
			6.95	9.00	12.280	118.13	0.08	0.008465	0.000006	1	3.04E+10	3097.8	94258168596401	257578309	797913677997	2.92E+17	0.008477	-0.000012	-372636.7	4.564		117.96	0.171
			7.00	10.00	12.375	118.37	0.08	0.008448	0.000006	1	3.07E+10	3847.8	118042047121273	259161004	997210532483	4.54E+17	0.008477	-0.000029	-888679.1	25.744		117.97	0.405
			7.08	12.00	12.500	118.69	0.08	0.008425	0.000006	1	3.10E+10	5584.9	173177161759360	261253356	1459071811294	9.67E+17	0.008476	-0.000050	-1560280.2	78.511		117.99	0.705
			7.15	14.00	12.595	118.93	0.08	0.008408	0.000006	1	3.13E+10	7637.8	238766203779957	262851053	2007595850059	1.82E+18	0.008474	-0.000066	-2061073.0	135.888		118.01	0.925
			7.20	16.00	12.630	119.02	0.08	0.008402	0.000006	1	3.14E+10	10006.5	313751843604660	263441317	2636118924914	3.14E+18	0.008472	-0.000070	-2210276.0	155.807		118.03	0.990
			7.26	18.00	12.612	118.97	0.08	0.008405	0.000006	1	3.13E+10	12691.0	397313563376623	263137642	3339481741856	5.04E+18	0.008470	-0.000065	-2044604.9	133.531		118.06	0.917
			7.30	20.00	12.545	118.80	0.08	0.008417	0.000006	1	3.11E+10	15691.4	488438429590110	262009350	4111284884279	7.66E+18	0.008468	-0.000051	-1589973.7	81.214		118.09	0.717
			7.34	22.00	12.432	118.52	0.08	0.008438	0.000006	1	3.08E+10	19007.6	585963527895071	260113725	4944125705453	1.11E+19	0.008466	-0.000028	-871475.9	24.636		118.12	0.396
			7.38	24.00	12.380	118.38	0.08	0.008447	0.000006	1	3.07E+10	22639.6	694822968871512	259244482	5869183229096	1.57E+19	0.008463	-0.000016	-497418.7	8.062		118.16	0.227
			7.41	26.00	12.210	117.95	0.08	0.008478	0.000006	1	3.02E+10	26587.4	804137062390975	256416246	6817443936990	2.14E+19	0.008460	0.000018	532521.2	9.376		118.20	-0.245
			7.45	28.00	12.125	117.74	0.08	0.008494	0.000006	1	3.00E+10	30851.1	926274448672837	255009874	7867329877120	2.86E+19	0.008457	0.000036	1088812.5	39.485		118.24	-0.505
			7.48	30.00	11.922	117.22	0.08	0.008531	0.000006	1	2.95E+10	35430.6	1045244674194240	251671954	8916882251216	3.70E+19	0.008454	0.000077	2271017.1	174.824		118.29	-1.067
			sums >>>>>>>					29	8.09E+11		6152813503205580	6856262305	52076778904748	1.34E+20			2.35E-06	1089.831
Derived parameters		esd					Chi^2(/n-2) >>>>>		40.364
Denom	7.03E+31					esd of fit >>>>>>		6.353
a	0.008480	0.000009					Ci'measured vs Log(f) (with expectation value)
C1' /pF	117.929	0.122
b	-7.27E-10	6.82E-10
L1 /nH	-0.727	0.682

KT	11.99	< Transformer const. = ( N R0 / k Ri) + (1/N) -1
L2 / nH	900.00	20.00
L1(true)	74.31	1.8	< corrected for L2

				C1'm obs-calc vs Log(f)

## Sheet3
Stray capacitance of voltage sampling network				Predicted C1 = C1' + C1s = C2 [( (1+Cx/C2) N R0 / k Ri ) + (1/N) -1]
		esd	Deriv dC1/dpar	(der*esd)^2
N =	12
R0 =	49.89	0.06	2.74	0.03
C2+Cx /pF =	10.9	0.2	11.64	5.42
Cx /pF	0.3	0.05
C2 /pF	10.6
1+Cx/C2	1.03
Ri =	49.72	0.06	-2.75	0.03
k =	0.959	0.0112	-142.71	2.55
				8.03	<<< sum
KT	11.995	< Transformer const. = [ (1+Cx/C2) N R0 / k Ri ] + (1/N) -1

C1' =	117.929	1.186
C1 pred =	127.142	2.833
C1s =	9.213	3.071

C1 prediction for bridge evaluation (see sheet 5)
R0 =	50
C1 pred =	127.169

Self capacitance of transformer secondary				Ci' = Ci apparent + ( Csh * Ro / Ri ) + [ L2 / ( N Ro Rv ) ] + ( LH / Ri^2 ) - Cis - [ L0 / ( R0 Ri ) ] + (C0 Ro / Ri) - [(N-1)Cx R0/Ri] -Cx C2/(N(Cx+C2)
parameter		esd +/-
Csh / pF	0.00	0.00	< shield protrusion capacitance. Set to 0 for gen. side earth.
L2 / nH	900.00	20.00	< inductance of upper voltage sampling arm
L2 fit	1		< =1 when L2 is included in fit on sheet 1, 0 otherwise.
RV0	1475.27		< optimal Rv
LH / nH	15.00	2.00	< inductance of sec. load resistor
Cis / pF	0.40	0.10	< strays accross sec winding
L0 / nH	13.00	3.00	< inductance of pri. transmission line
C0 / pF	3.50	0.50	< capacitance of pri. transmission line
Cneut /pF	0.00	0.00

Series for Ci'		esd +/-	Variance
Ci apparent	6.05	0.08	0.006	< from fit on sheet 1
Corr for Csh	0.00	0.00	0.000
Corr for Cx	-3.34	0.55	0.305
Corr for L2	0.00	0.02	0.001	< inclusion depends on whether L2 is included in fit on sheet 1.
Corr for LH	6.07	0.81	0.655
Corr for Cis	-0.40	0.10	0.010
Corr for L0	-5.24	1.21	1.463
Corr for C0	3.51	0.50	0.252
Ci' / pF	6.65	1.64	2.69	Only valid if parameters have been filled in correctly!

lw / mm	228.00	< length of secondary winding wire
Ci' pred / pF	7.65	< predicted Ci' if phase velocity = c

## Sheet4
Data reduction for 8-48 reference capacitor with 82pF in parallel. Data for dial readings 0 and 18 are excluded from fit (end effects).
		Dial	C /pF	esd C	fit	w	wy	wx	wx^2	wxy	y calc	obs-calc	we	we^2
		x	y					C calc	e
		0	89.0	0.20	0	0.000	0.00	0.00	0.00	0.00	86.92	2.08	0.00	0.000
		1	90.0	0.20	1	25.000	2250.00	25.00	25.00	2250.00	89.46	0.54	13.48	7.269
		2	92.0	0.20	1	25.000	2300.00	50.00	100.00	4600.00	92.00	0.00	-0.06	0.000
		3	94.5	0.20	1	25.000	2362.50	75.00	225.00	7087.50	94.54	-0.04	-1.10	0.049
		4	97.0	0.20	1	25.000	2425.00	100.00	400.00	9700.00	97.09	-0.09	-2.14	0.184
		5	99.5	0.20	1	25.000	2487.50	125.00	625.00	12437.50	99.63	-0.13	-3.19	0.406
		6	102.0	0.20	1	25.000	2550.00	150.00	900.00	15300.00	102.17	-0.17	-4.23	0.715
		7	104.5	0.20	1	25.000	2612.50	175.00	1225.00	18287.50	104.71	-0.21	-5.27	1.111
		8	107.0	0.20	1	25.000	2675.00	200.00	1600.00	21400.00	107.25	-0.25	-6.31	1.593
		9	110.0	0.20	1	25.000	2750.00	225.00	2025.00	24750.00	109.79	0.21	5.15	1.060
		10	112.0	0.20	1	25.000	2800.00	250.00	2500.00	28000.00	112.34	-0.34	-8.39	2.819
		11	115.0	0.20	1	25.000	2875.00	275.00	3025.00	31625.00	114.88	0.12	3.06	0.375
		12	117.5	0.20	1	25.000	2937.50	300.00	3600.00	35250.00	117.42	0.08	2.02	0.164
		13	120.0	0.20	1	25.000	3000.00	325.00	4225.00	39000.00	119.96	0.04	0.98	0.038
		14	122.5	0.20	1	25.000	3062.50	350.00	4900.00	42875.00	122.50	0.00	-0.06	0.000
		15	125.0	0.20	1	25.000	3125.00	375.00	5625.00	46875.00	125.04	-0.04	-1.10	0.049
		16	128.0	0.20	1	25.000	3200.00	400.00	6400.00	51200.00	127.59	0.41	10.36	4.289
		17	130.0	0.20	1	25.000	3250.00	425.00	7225.00	55250.00	130.13	-0.13	-3.19	0.406
		18	131.0	0.20	0	0.000	0.00	0.00	0.00	0.00	132.67	-1.67	0.00	0.000
		sums >>>>>>>>>			17	425.000	46662.50	3825.00	44625.00	445887.50			-1.49E-11	20.527
D	4335000
a	86.919118		Capacitance vs dial reading
esd a	0.014087
b	2.541667
esd b	0.000134
Chi^2/n-2)	1.368464
esd fit	1.169814

## Sheet5
Bridge performance evaluation				FT50-61, 12 turns 0.9mm wire + 1 turn shift comp with 3.34 Ohm load. Lsec = 9.25uH @ 1.5915MHz. 40pF variable ref cap with padding. No shield. L-bal coil.
N	12
R0	50
C2 /pF	10.9
Li /uH	9.627
Rv0	1472.03	<<< Guide value (estimated value for zero phase error at 0Hz)
Rvset	1486.00	<<< Selected value
C1s /pF	9.213
L1 /nH	-0.727
C1 pred	127.169								| << - - show hidden columns for details of fit - - >> |
Cv set 0	138.069	<<< Guide value (est. value for zero magnitude error at 0Hz)
C1 set	127.310
Cv set	138.210	<<< Selected value
			x						y				| - - - Determining the gradient of Phi vs f: y=phi, x=f
			f	Rv	C1'm	C1'0	Cv	Xcv	Xcvset	U	Rbal	Xbal	|Zbal|	(|Zbal|/R0)-1	Phi				esd phi	fit	weight					Phi calc	obs-calc
		Log(f)	/MHz	/Ohms	/pF	/pF	/pF				Ohms	Ohms	Ohms	Magn. err.	/deg.					w	wx	wy	wxy	wx^2	Y calc	e	we	we^2
		6.20	1.6	1483.2	117.55	117.55	137.66	-726.4	-723.5	3.67E-07	50.143	0.116	50.143	0.0029	0.133				0.0500	1	400.00	636.6	53.1	84.5	1013.2	-0.152	0.285	113.838	32.397905
		6.23	1.7	1482.6	117.56	117.56	137.67	-680.0	-677.4	3.76E-07	50.141	0.117	50.141	0.0028	0.134				0.0500	1	400.00	680.0	53.6	91.1	1156.0	-0.152	0.286	114.541	32.798900
		6.26	1.8	1483.2	117.55	117.55	137.66	-642.3	-639.7	3.84E-07	50.153	0.107	50.153	0.0031	0.122				0.0500	1	400.00	720.0	49.0	88.2	1296.0	-0.153	0.275	110.068	30.287667
		6.28	1.9	1485.2	117.56	117.56	137.67	-608.4	-606.1	3.90E-07	50.163	0.078	50.163	0.0033	0.089				0.0500	1	400.00	760.0	35.6	67.6	1444.0	-0.153	0.242	96.821	23.435688
Calibration freq. >>>		6.30	2.0	1485.8	117.55	117.55	137.66	-578.1	-575.8	3.95E-07	50.173	0.070	50.173	0.0035	0.080				0.0500	1	400.00	800.0	31.9	63.7	1600.0	-0.153	0.233	93.259	21.743296
		6.34	2.2	1488.8	117.55	117.55	137.66	-525.5	-523.4	4.04E-07	50.188	0.033	50.188	0.0038	0.038				0.0500	1	400.00	880.0	15.1	33.3	1936.0	-0.154	0.192	76.820	14.753328
		6.38	2.4	1491.7	117.56	117.56	137.67	-481.7	-479.8	4.10E-07	50.195	0.001	50.195	0.0039	0.001				0.0500	1	400.00	960.0	0.4	1.1	2304.0	-0.155	0.156	62.443	9.747715
		6.41	2.6	1494.4	117.60	117.60	137.71	-444.5	-442.9	4.15E-07	50.189	-0.028	50.189	0.0038	-0.032				0.0500	1	400.00	1040.0	-12.7	-32.9	2704.0	-0.156	0.124	49.648	6.162358
		6.45	2.8	1498.6	117.61	117.61	137.72	-412.7	-411.3	4.19E-07	50.194	-0.063	50.194	0.0039	-0.072				0.0500	1	400.00	1120.0	-28.9	-81.0	3136.0	-0.157	0.084	33.696	2.838539
		6.48	3.0	1501.9	117.65	117.65	137.76	-385.1	-383.8	4.22E-07	50.186	-0.090	50.186	0.0037	-0.102				0.0500	1	400.00	1200.0	-40.9	-122.7	3600.0	-0.157	0.055	22.010	1.211137
		6.54	3.5	1512.5	117.70	117.70	137.81	-330.0	-329.0	4.26E-07	50.178	-0.156	50.178	0.0036	-0.178				0.0500	1	400.00	1400.0	-71.1	-248.8	4900.0	-0.159	-0.019	-7.417	0.137521
		6.60	4.0	1523.6	117.76	117.76	137.88	-288.6	-287.9	4.27E-07	50.161	-0.209	50.161	0.0032	-0.239				0.0500	1	400.00	1600.0	-95.6	-382.3	6400.0	-0.161	-0.078	-31.138	2.424014
		6.65	4.5	1535.6	117.56	117.56	137.67	-256.9	-255.9	4.28E-07	50.235	-0.240	50.236	0.0047	-0.274				0.0500	1	400.00	1800.0	-109.4	-492.4	8100.0	-0.163	-0.111	-44.220	4.888535
		6.70	5.0	1547.8	117.61	117.61	137.72	-231.1	-230.3	4.27E-07	50.218	-0.278	50.219	0.0044	-0.317				0.0500	1	400.00	2000.0	-126.8	-634.2	10000.0	-0.165	-0.152	-60.879	9.265725
		6.78	6.0	1577.3	117.75	117.75	137.86	-192.4	-191.9	4.21E-07	50.169	-0.354	50.170	0.0034	-0.404				0.0500	1	400.00	2400.0	-161.8	-970.7	14400.0	-0.169	-0.236	-94.286	22.224686
		6.85	7.0	1617.6	117.90	117.90	138.02	-164.7	-164.5	4.12E-07	50.115	-0.439	50.117	0.0023	-0.502				0.0500	1	400.00	2800.0	-200.9	-1406.1	19600.0	-0.173	-0.330	-131.851	43.461495
		6.90	8.0	1657.6	118.21	118.21	138.32	-143.8	-143.9	4.03E-07	50.003	-0.501	50.006	0.0001	-0.574				0.0500	1	400.00	3200.0	-229.5	-1835.9	25600.0	-0.176	-0.397	-158.945	63.159179
		6.95	9.0	1693.5	118.13	118.13	138.24	-127.9	-127.9	3.95E-07	50.027	-0.525	50.030	0.0006	-0.602				0.0500	1	400.00	3600.0	-240.6	-2165.8	32400.0	-0.180	-0.421	-168.588	71.054961
		7.00	10.0	1733.4	118.37	118.37	138.49	-114.9	-115.2	3.86E-07	49.937	-0.556	49.940	-0.0012	-0.638				0.0500	1	400.00	4000.0	-255.2	-2552.0	40000.0	-0.184	-0.454	-181.611	82.456452
		7.08	12.0	1801.3	118.69	118.69	138.80	-95.6	-96.0	3.71E-07	49.816	-0.573	49.820	-0.0036	-0.658				0.0500	0	0.00	0.0	0.0	0.0	0.0	-0.192	-0.467	0.000	0.000000
		7.15	14.0	1859.0	118.93	118.93	139.05	-81.8	-82.3	3.59E-07	49.724	-0.564	49.728	-0.0054	-0.650				0.0500	0	0.00	0.0	0.0	0.0	0.0	-0.199	-0.451	0.000	0.000000
		7.20	16.0	1880.2	119.02	119.02	139.13	-71.5	-72.0	3.55E-07	49.687	-0.516	49.690	-0.0062	-0.595				0.0500	0	0.00	0.0	0.0	0.0	0.0	-0.207	-0.388	0.000	0.000000
		7.26	18.0	1859.4	118.97	118.98	139.09	-63.6	-64.0	3.59E-07	49.699	-0.440	49.701	-0.0060	-0.507				0.0500	1	400.00	7200.0	-202.8	-3650.7	129600.0	-0.214	-0.293	-117.043	34.247556
		7.30	20.0	1828.7	118.80	118.81	138.92	-57.3	-57.6	3.66E-07	49.757	-0.369	49.758	-0.0048	-0.425				0.0500	1	400.00	8000.0	-169.9	-3397.7	160000.0	-0.222	-0.203	-81.060	16.426803
		7.34	22.0	1762.5	118.52	118.52	138.63	-52.2	-52.3	3.80E-07	49.856	-0.280	49.857	-0.0029	-0.321				0.0500	1	400.00	8800.0	-128.6	-2828.7	193600.0	-0.230	-0.092	-36.707	3.368497
Calibration freq. >>>		7.38	24.0	1680.8	118.38	118.39	138.50	-47.9	-48.0	3.99E-07	49.901	-0.190	49.901	-0.0020	-0.218				0.0500	1	400.00	9600.0	-87.0	-2089.2	230400.0	-0.237	0.020	7.869	0.154800
		7.41	26.0	1583.2	117.95	117.95	138.07	-44.3	-44.3	4.25E-07	50.054	-0.090	50.054	0.0011	-0.103				0.0500	1	400.00	10400.0	-41.2	-1072.2	270400.0	-0.245	0.142	56.727	8.044843
		7.45	28.0	1476.3	117.74	117.74	137.85	-41.2	-41.1	4.57E-07	50.130	0.013	50.130	0.0026	0.015				0.0500	1	400.00	11200.0	5.8	163.1	313600.0	-0.253	0.267	106.839	28.536565
		7.48	30.0	1347.2	117.22	117.22	137.33	-38.6	-38.4	5.02E-07	50.315	0.143	50.315	0.0063	0.163				0.0500	1	400.00	12000.0	65.1	1953.2	360000.0	-0.260	0.423	169.166	71.542803
									Sums >>>>		26	1.04E+04	98796.6	-1893.4	-21417.5	1839189.2		-1.31E+00	0.00E+00	636.770966
									Parameter		esd					Chi^2/n-2 >>>>>>>>>>>			26.532124
phase error vs. f									D	9.37E+09					esd of an obs of unit weight >>>			5.151
									a	-0.145869	0.072178	<<< phase intercept at 0Hz
									b	-0.003809	0.005428	<<< gradient [degrees / MHz]
									tp	0.0106	0.015077	ns <<< time delay

									Phase error, obs-calc

Magnitude error vs f