## Sheet1
FT50-61, 12 turns 0.9mm wire. Gen side shield earth. Capacitor from load terminal to 1-turn winding loaded with 4.26 Ohms.. Lsec = 9.25uH @ 1.5915MHz.. L-bal coil.					D. W. Knight. June 2007.			Version	1.01
		esd +/-					Show columns N-Q to see deatails of fitting procedure.					See sheet 2 for C1'm vs frequency
N =	12.00					Li' = N R0 C2 Rv
R0 =	49.89	0.06					from fit	from fit
C2+Cx/pF	10.30	0.20	< Adjusted to reproduce Li as measured elsewhere.					Ci =	Li =
L2 / nH	900.00	20.00					Li computed from Rv
L2 fit	1	< 1 to include L2 in fit, 0 to exclude it.
1/(N R0 C2)	1.62E+08					1.039	8.498
			Log (f)	f	x	Rv	esd Rv	Li'	Corr(L2)	y	esd y	fit	weight	w x	w x^2	w y	w x y	y calc	obs-calc	w e	w e^2	Li' calc	Li'					Computed	Computed
				/MHz	-(2 pi f)^2	Ohms	Ohms	/ uH					w					e			/ uH	obs-calc					Li /uH	Ci /pF
			6.20	1.59	-100.000	1440.0	0.86	8.8796	-104.4	112513.01	67.6	1	2.19E-04	-0.021902150	2.190217	24.642751	-2464.276858035	117571.42	-5058.41	-1.107900	5604.21	8.4979	0.3817					8.871	50.580
			6.23	1.70	-114.093	1438.8	0.86	8.8722	-119.2	112592.13	67.6	1	2.19E-04	-0.024947090	2.846279	24.618996	-2808.845896248	117556.78	-4964.65	-1.085552	5389.39	8.4979	0.3743					8.863	43.509
			6.26	1.80	-127.910	1436.3	0.86	8.8568	-133.9	112773.64	67.7	1	2.18E-04	-0.027871254	3.565014	24.573068	-3143.142864605	117542.41	-4768.77	-1.039102	4955.24	8.4979	0.3589					8.846	37.275
			6.28	1.90	-142.517	1436.0	0.86	8.8550	-149.2	112781.91	67.8	1	2.18E-04	-0.031041112	4.423889	24.564604	-3500.875868691	117527.23	-4745.32	-1.033561	4904.58	8.4979	0.3571					8.843	33.289
Calibration freq. >>>			6.30	2.00	-157.914	1435.7	0.86	8.8531	-165.3	112789.35	67.8	1	2.18E-04	-0.034380214	5.429106	24.555963	-3877.722176536	117511.23	-4721.87	-1.028024	4854.20	8.4979	0.3552					8.840	29.894
			6.34	2.20	-191.076	1431.4	0.86	8.8266	-200.7	113093.35	68.0	1	2.16E-04	-0.041351243	7.901211	24.474879	-4676.550733848	117476.76	-4383.41	-0.948627	4158.22	8.4978	0.3288					8.811	22.930
			6.38	2.40	-227.396	1428.2	0.86	8.8069	-239.4	113308.52	68.1	1	2.15E-04	-0.048991612	11.140481	24.411927	-5551.166863021	117439.01	-4130.49	-0.889900	3675.72	8.4977	0.3091					8.789	18.151
			6.41	2.60	-266.874	1424.1	0.85	8.7816	-281.7	113593.06	68.3	1	2.14E-04	-0.057167454	15.256513	24.332919	-6493.826046005	117397.97	-3804.91	-0.815056	3101.22	8.4976	0.2839					8.760	14.241
			6.45	2.80	-309.511	1420.4	0.85	8.7588	-327.6	113843.83	68.5	1	2.13E-04	-0.065956648	20.414294	24.260083	-7508.757580426	117353.66	-3509.82	-0.747942	2625.15	8.4975	0.2612					8.734	11.321
			6.48	3.00	-355.306	1416.7	0.85	8.7359	-377.0	114092.56	68.7	1	2.12E-04	-0.075321591	26.762195	24.186586	-8593.633423620	117306.06	-3213.49	-0.681231	2189.13	8.4974	0.2385					8.708	9.023
			6.54	3.50	-483.611	1406.0	0.84	8.6700	-517.1	114823.65	69.2	1	2.09E-04	-0.100978268	48.834162	23.975267	-11594.693425906	117172.70	-2349.04	-0.490482	1152.16	8.4969	0.1730					8.632	4.828
			6.60	4.00	-631.655	1394.8	0.84	8.6009	-680.8	115586.10	69.8	1	2.05E-04	-0.129797117	81.986956	23.751495	-15002.743106435	117018.82	-1432.71	-0.294405	421.80	8.4962	0.1047					8.553	2.230
			6.65	4.50	-799.438	1383.4	0.83	8.5306	-868.7	116356.27	70.3	1	2.02E-04	-0.161600149	129.189292	23.520513	-18803.191081979	116844.42	-488.15	-0.098676	48.17	8.4952	0.0354					8.471	0.563
			6.70	5.00	-986.960	1370.8	0.82	8.4529	-1082.4	117220.14	71.0	1	1.98E-04	-0.195888701	193.334398	23.265472	-22962.100480664	116649.51	570.63	0.113256	64.63	8.4939	-0.0410					8.380	-0.635
			6.78	6.00	-1421.223	1348.3	0.81	8.3142	-1584.6	118692.08	72.2	1	1.92E-04	-0.272895739	387.845711	22.790627	-32390.563589655	116198.13	2493.95	0.478875	1194.29	8.4902	-0.1760					8.213	-1.830
			6.85	7.00	-1934.442	1331.8	0.80	8.2124	-2183.6	119583.28	73.1	1	1.87E-04	-0.362405923	701.053405	22.403193	-43337.687812855	115664.69	3918.59	0.734123	2876.73	8.4855	-0.2731					8.079	-2.115
			6.90	8.00	-2526.619	1312.6	0.79	8.0940	-2893.7	120654.26	74.1	1	1.82E-04	-0.459796816	1161.731245	21.956794	-55476.445830677	115049.18	5605.09	1.020020	5717.30	8.4787	-0.3847					7.926	-2.324
			6.95	9.00	-3197.752	1296.7	0.78	7.9960	-3707.3	121355.65	75.0	1	1.78E-04	-0.567917467	1816.059116	21.552639	-68919.991013752	114351.60	7004.05	1.243912	8712.42	8.4704	-0.4744					7.789	-2.310	Ci computed from Rv
			7.00	10.00	-3947.842	1284.3	0.77	7.9195	-4621.1	121649.34	75.8	1	1.74E-04	-0.687787297	2715.275414	21.193573	-83668.872396755	113571.95	8077.40	1.407232	11366.77	8.4607	-0.5412					7.670	-2.177
			7.08	12.00	-5684.892	1265.5	0.76	7.8036	-6753.2	121393.07	76.9	1	1.69E-04	-0.961629940	5466.762482	20.534288	-116735.210524746	111766.44	9626.63	1.628395	15675.95	8.4374	-0.6338					7.460	-1.842
			7.15	14.00	-7737.770	1259.2	0.76	7.7647	-9237.8	119549.57	77.3	1	1.67E-04	-1.295885668	10027.265051	20.021606	-154922.579949656	109632.67	9916.91	1.660838	16470.37	8.4125	-0.6478					7.308	-1.436
			7.20	16.00	-10106.475	1260.6	0.76	7.7734	-12052.3	116592.05	77.2	1	1.68E-04	-1.696351145	17144.130278	19.569736	-197781.049686550	107170.62	9421.43	1.581368	14898.74	8.3876	-0.6143					7.187	-1.085
			7.26	18.00	-12791.007	1264.8	0.76	7.7993	-15203.0	113014.12	76.9	1	1.69E-04	-2.161274400	27644.876639	19.095800	-244254.518211637	104380.29	8633.83	1.458843	12595.39	8.3624	-0.5631					7.067	-0.824
			7.30	20.00	-15791.367	1282.6	0.77	7.9090	-18508.7	107929.04	75.9	1	1.74E-04	-2.743870736	43329.469903	18.753497	-296143.347254240	101261.69	6667.35	1.158503	7724.15	8.3493	-0.4403					7.000	-0.555
			7.34	22.00	-19107.554	1306.2	0.78	8.0546	-21990.9	102162.41	74.5	1	1.80E-04	-3.443387557	65794.714112	18.410770	-351784.783585565	97814.82	4347.59	0.783483	3406.26	8.3468	-0.2923					6.944	-0.339
			7.38	24.00	-22739.569	1338.8	0.80	8.2556	-25533.7	95596.44	72.7	1	1.89E-04	-4.305019036	97894.275427	18.098166	-411544.484216473	94039.68	1556.76	0.294723	458.81	8.3631	-0.1075					6.908	-0.152
			7.41	26.00	-26687.410	1386.2	0.83	8.5479	-28942.0	88046.24	70.2	1	2.03E-04	-5.416511509	144552.665046	17.869978	-476903.447761968	89936.26	-1890.03	-0.383602	725.02	8.4120	0.1359					6.910	0.026
			7.45	28.00	-30951.079	1439.5	0.86	8.8765	-32323.0	80333.51	67.6	1	2.19E-04	-6.774240118	209670.043766	17.582536	-544198.482357776	85504.57	-5171.06	-1.131786	5852.54	8.4870	0.3896					6.905	0.162
			7.48	30.00	-35530.576	1487.8	0.89	9.1744	-35900.9	73098.33	65.4	1	2.34E-04	-8.307164366	295158.333549	17.090628	-607239.848472902	80744.61	-7646.28	-1.787724	13669.43	8.5730	0.6014					6.853	0.244
Derived parameters		esd +/-	sums >>>>>>>>>>>>					29	5.76E-03	-40.473332317	924017.775153	636.058354	-3802282.839071230			-2.19E-13	164487.98
D	3.69E+03
a	117675.36	1235.68
Li / uH	8.497955	0.0892		y obs - y calc
b = Ci / pF	1.039408	0.0976
chi^2/(n-2)	6092.148
esd of fit	78.05
Rv0	1378.11
				Incremental error analysis
X-over /MHz	53.55					R0	49.83	49.89	49.95	err contrib.
				Li	8.994662	8.497955	9.016323	0.010831
				Ci	6.173432	1.039408	6.158601	-0.007416

				C2	10.1	10.3	10.5
				Li	8.828914	8.497955	9.182070	0.176578
				Ci	6.289328	1.039408	6.047431	-0.120949

				Accuracy
				esd Li	0.198141
				esd Ci	0.155583

				Li' vs Log(f)

## Sheet2
Variation of C1'm (nominal capacitance, measured @ 1.5915MHz) with frequency					To see details of fitting calculation show columns N-P
			Log(f)	f	Dial	Ci'm	esd C1'm	y	esd y	fit	weight	x	wx	wy	wxy	wx^2	y calc	e	we	we^2		<Ci'm>	obs-calc
				/MHz		/pF		1/C1'm			w					obs-calc
			6.20	1.59	11.840	117.01	0.63	0.008546	0.000046	1	4.80E+08	0.0	34320	4101438	293.31	2.45E+00	0.008604	-0.000058	-27787.7	1.609		116.23	0.787
			6.23	1.70	11.820	116.96	0.13	0.008550	0.000009	1	1.20E+10	14.1	168789486623	102402385	1443118606	2.38E+12	0.008604	-0.000054	-649070.1	35.175		116.22	0.737
			6.26	1.80	11.810	116.94	0.10	0.008552	0.000007	1	1.87E+10	27.9	521863466233	159899439	4462805033	1.46E+13	0.008604	-0.000052	-978649.2	51.222		116.22	0.711
			6.28	1.90	11.805	116.92	0.10	0.008553	0.000007	1	1.87E+10	42.5	794640373423	159847313	6796242187	3.38E+13	0.008604	-0.000051	-960963.9	49.409		116.22	0.699
			6.30	2.00	11.805	116.92	0.08	0.008553	0.000006	1	2.92E+10	57.9	1691251658187	249761426	14464600910	9.79E+13	0.008604	-0.000051	-1501690.2	77.221		116.22	0.699
			6.34	2.20	11.790	116.89	0.08	0.008555	0.000006	1	2.92E+10	91.1	2656209804724	249517188	22724912948	2.42E+14	0.008604	-0.000049	-1418769.1	69.018		116.22	0.661
			6.38	2.40	11.765	116.82	0.08	0.008560	0.000006	1	2.91E+10	127.4	3707410763397	249110479	31735600195	4.72E+14	0.008604	-0.000044	-1280697.0	56.361		116.22	0.598
			6.41	2.60	11.765	116.82	0.08	0.008560	0.000006	1	2.91E+10	166.9	4856293552381	249110479	41570087709	8.10E+14	0.008604	-0.000044	-1281167.5	56.402		116.22	0.598
			6.45	2.80	11.745	116.77	0.08	0.008564	0.000006	1	2.91E+10	209.5	6086481651160	248785430	52123232947	1.28E+15	0.008604	-0.000040	-1171190.9	47.217		116.22	0.547
			6.48	3.00	11.740	116.76	0.08	0.008565	0.000006	1	2.90E+10	255.3	7413639351292	248704212	63495617419	1.89E+15	0.008604	-0.000039	-1144159.0	45.082		116.22	0.535
			6.54	3.50	11.700	116.66	0.08	0.008572	0.000006	1	2.89E+10	383.6	11100643722769	248055104	95156571026	4.26E+15	0.008604	-0.000032	-925707.0	29.613		116.22	0.434
			6.60	4.00	11.655	116.54	0.08	0.008581	0.000006	1	2.88E+10	531.7	15324388472612	247326208	131492136452	8.15E+15	0.008604	-0.000024	-681340.6	16.106		116.22	0.320
			6.65	4.50	11.600	116.40	0.08	0.008591	0.000006	1	2.87E+10	699.4	20064009387926	246437277	172367585209	1.40E+16	0.008604	-0.000013	-384448.7	5.152		116.22	0.181
			6.70	5.00	11.560	116.30	0.08	0.008598	0.000006	1	2.86E+10	887.0	25354488306992	245792121	218007887536	2.25E+16	0.008604	-0.000006	-170626.3	1.018		116.22	0.081
			6.78	6.00	11.450	116.02	0.08	0.008619	0.000006	1	2.83E+10	1321.2	37406372063395	244023750	322409798789	4.94E+16	0.008605	0.000015	412598.5	6.013		116.22	-0.197
			6.85	7.00	11.355	115.78	0.08	0.008637	0.000006	1	2.81E+10	1834.4	51505598753098	242503363	444858466391	9.45E+16	0.008605	0.000032	907960.6	29.362		116.21	-0.435
			6.90	8.00	11.310	115.67	0.08	0.008646	0.000006	1	2.80E+10	2426.6	67863294829857	241785389	586720953820	1.65E+17	0.008605	0.000041	1136447.5	46.181		116.21	-0.546
			6.95	9.00	11.230	115.46	0.08	0.008661	0.000006	1	2.78E+10	3097.8	86024759619559	240512494	745048016267	2.66E+17	0.008605	0.000056	1543652.1	85.807		116.21	-0.746
			7.00	10.00	11.185	115.35	0.08	0.008669	0.000006	1	2.77E+10	3847.8	106432036934451	239798457	922706516696	4.10E+17	0.008606	0.000064	1766588.4	112.828		116.20	-0.856
			7.08	12.00	11.115	115.17	0.08	0.008683	0.000006	1	2.75E+10	5584.9	153528285926828	238690544	1333060944638	8.57E+17	0.008606	0.000077	2104323.2	161.084		116.19	-1.024
			7.15	14.00	11.080	115.08	0.08	0.008690	0.000006	1	2.74E+10	7637.8	209313792183041	238137870	1818842245743	1.60E+18	0.008607	0.000082	2258728.5	186.164		116.18	-1.102
			7.20	16.00	11.110	115.16	0.08	0.008684	0.000006	1	2.75E+10	10006.5	274955887708749	238611539	2387660374786	2.75E+18	0.008608	0.000076	2079966.0	157.446		116.17	-1.013
			7.26	18.00	11.160	115.28	0.08	0.008674	0.000006	1	2.76E+10	12691.0	350262819960136	239402381	3038257368916	4.45E+18	0.008609	0.000065	1794622.6	116.694		116.15	-0.871
			7.30	20.00	11.240	115.49	0.08	0.008659	0.000006	1	2.78E+10	15691.4	436134045974139	240671361	3776462656953	6.84E+18	0.008610	0.000049	1348677.6	65.442		116.14	-0.651
			7.34	22.00	11.340	115.74	0.08	0.008640	0.000006	1	2.80E+10	19007.6	532972073331469	242263881	4604843825120	1.01E+19	0.008612	0.000028	789329.2	22.220		116.12	-0.378
			7.38	24.00	11.390	115.87	0.08	0.008630	0.000006	1	2.82E+10	22639.6	637606482946059	243062768	5502836206423	1.44E+19	0.008613	0.000017	484028.3	8.319		116.10	-0.231
			7.41	26.00	11.545	116.26	0.08	0.008601	0.000006	1	2.85E+10	26587.4	759026735861519	245550478	6528551302214	2.02E+19	0.008615	-0.000014	-390394.7	5.339		116.08	0.185
			7.45	28.00	11.710	116.68	0.08	0.008570	0.000006	1	2.90E+10	30851.1	893524281051773	248217275	7657770849704	2.76E+19	0.008617	-0.000046	-1341984.5	62.181		116.05	0.627
			7.48	30.00	11.875	117.10	0.08	0.008540	0.000006	1	2.94E+10	35430.6	1040990399368930	250903310	8889648770836	3.69E+19	0.008619	-0.000079	-2318276.0	182.920		116.03	1.072
			sums >>>>>>>					29	7.60E+11		5737286976545040	6542985359	49415518695768	1.27E+20			-9.92E-07	1788.604
Derived parameters		esd					Chi^2(/n-2) >>>>>		66.245
Denom	6.34E+31					esd of fit >>>>>>		8.139
a	0.008604	0.000012					Ci'measured vs Log(f) (with expectation value)
C1' /pF	116.225	0.155
b	4.10E-10	8.91E-10
L1 /nH	0.410	0.891

KT	12.68	< Transformer const. = ( N R0 / k Ri) + (1/N) -1
L2 / nH	900.00	20.00
L1(true)	71.41	1.81	< corrected for L2

				C1'm obs-calc vs Log(f)

## Sheet3
Stray capacitance of voltage sampling network				Predicted C1 = C1' + C1s = C2 [( (1+Cx/C2) N R0 / k Ri ) + (1/N) -1]
		esd	Deriv dC1/dpar	(der*esd)^2
N =	12
R0 =	49.89	0.06	2.81	0.03
C2+Cx /pF =	10.3	0.2	12.68	6.43
Cx /pF	0	0
C2 /pF	10.3
1+Cx/C2	1
Ri =	49.77	0.06	-2.81	0.03
k =	0.885	0.0200	-158.19	10.01	<<<<< 0.959 x 12 / 13
				16.49	<<< sum
KT	12.675	< Transformer const. = [ (1+Cx/C2) N R0 / k Ri ] + (1/N) -1

C1' =	116.225	1.173
C1 pred =	130.556	4.061
C1s =	14.331	4.227

C1 prediction for bridge evaluation (see sheet 5)
R0 =	50
C1 pred =	130.865

Self capacitance of transformer secondary				Ci' = Ci apparent + ( Csh * Ro / Ri ) + [ L2 / ( N Ro Rv ) ] + ( LH / Ri^2 ) - Cis - [ L0 / ( R0 Ri ) ] + (C0 Ro / Ri) - [(N-1)Cx R0/Ri] -Cx C2/(N(Cx+C2)
parameter		esd +/-
Cqi / pF	3.40	0.20	< neutralising capacitor from gen port to 1 turn current injection winding.
L2 / nH	900.00	20.00	< inductance of upper voltage sampling arm
L2 fit	1		< =1 when L2 is included in fit on sheet 1, 0 otherwise.
RV0	1378.11		< optimal Rv
LH / nH	15.00	2.00	< inductance of sec. load resistor
Cis / pF	0.40	0.10	< strays accross sec winding
L0 / nH	13.00	3.00	< inductance of pri. transmission line
C0 / pF	3.50	0.50	< capacitance of pri. transmission line
Cn0 /pF	3.40	0.20	< neutralising capacitor across C0.

Series for Ci'		esd +/-	Variance
Ci apparent	1.04	0.10	0.010	< from fit on sheet 1
Corr for Cqi	3.41	0.20	0.040
Corr for Cx	0.00	0.00	0.000
Corr for L2	0.00	0.02	0.001	< inclusion depends on whether L2 is included in fit on sheet 1.
Corr for LH	6.06	0.81	0.652
Corr for Cis	-0.40	0.10	0.010
Corr for L0	-5.24	1.21	1.460
Corr for C0	6.92	0.54	0.291
Ci' / pF	11.78	1.57	2.46	Only valid if parameters have been filled in correctly!

lw / mm	228.00	< length of secondary winding wire
Ci' pred / pF	7.64	< predicted Ci' if phase velocity = c

## Sheet4
Data reduction for 8-48 reference capacitor with 82pF in parallel. Data for dial readings 0 and 18 are excluded from fit (end effects).
		Dial	C /pF	esd C	fit	w	wy	wx	wx^2	wxy	y calc	obs-calc	we	we^2
		x	y					C calc	e
		0	89.0	0.20	0	0.000	0.00	0.00	0.00	0.00	86.92	2.08	0.00	0.000
		1	90.0	0.20	1	25.000	2250.00	25.00	25.00	2250.00	89.46	0.54	13.48	7.269
		2	92.0	0.20	1	25.000	2300.00	50.00	100.00	4600.00	92.00	0.00	-0.06	0.000
		3	94.5	0.20	1	25.000	2362.50	75.00	225.00	7087.50	94.54	-0.04	-1.10	0.049
		4	97.0	0.20	1	25.000	2425.00	100.00	400.00	9700.00	97.09	-0.09	-2.14	0.184
		5	99.5	0.20	1	25.000	2487.50	125.00	625.00	12437.50	99.63	-0.13	-3.19	0.406
		6	102.0	0.20	1	25.000	2550.00	150.00	900.00	15300.00	102.17	-0.17	-4.23	0.715
		7	104.5	0.20	1	25.000	2612.50	175.00	1225.00	18287.50	104.71	-0.21	-5.27	1.111
		8	107.0	0.20	1	25.000	2675.00	200.00	1600.00	21400.00	107.25	-0.25	-6.31	1.593
		9	110.0	0.20	1	25.000	2750.00	225.00	2025.00	24750.00	109.79	0.21	5.15	1.060
		10	112.0	0.20	1	25.000	2800.00	250.00	2500.00	28000.00	112.34	-0.34	-8.39	2.819
		11	115.0	0.20	1	25.000	2875.00	275.00	3025.00	31625.00	114.88	0.12	3.06	0.375
		12	117.5	0.20	1	25.000	2937.50	300.00	3600.00	35250.00	117.42	0.08	2.02	0.164
		13	120.0	0.20	1	25.000	3000.00	325.00	4225.00	39000.00	119.96	0.04	0.98	0.038
		14	122.5	0.20	1	25.000	3062.50	350.00	4900.00	42875.00	122.50	0.00	-0.06	0.000
		15	125.0	0.20	1	25.000	3125.00	375.00	5625.00	46875.00	125.04	-0.04	-1.10	0.049
		16	128.0	0.20	1	25.000	3200.00	400.00	6400.00	51200.00	127.59	0.41	10.36	4.289
		17	130.0	0.20	1	25.000	3250.00	425.00	7225.00	55250.00	130.13	-0.13	-3.19	0.406
		18	131.0	0.20	0	0.000	0.00	0.00	0.00	0.00	132.67	-1.67	0.00	0.000
		sums >>>>>>>>>			17	425.000	46662.50	3825.00	44625.00	445887.50			-1.49E-11	20.527
D	4335000
a	86.919118		Capacitance vs dial reading
esd a	0.014087
b	2.541667
esd b	0.000134
Chi^2/n-2)	1.368464
esd fit	1.169814

## Sheet5
Bridge performance evaluation				FT50-61, 12 turns 0.9mm wire. Gen side shield earth. Capacitor from load terminal to 1-turn winding loaded with 4.26 Ohms.. Lsec = 9.25uH @ 1.5915MHz.. L-bal coil.
N	12
R0	50
C2 /pF	10.3
Li /uH	8.498
Rv0	1375.07	<<< Guide value (estimated value for zero phase error at 0Hz)
Rvset	1438.00	<<< Selected value
C1s /pF	14.331
L1 /nH	0.410
C1 pred	130.865								| << - - show hidden columns for details of fit - - >> |
Cv set 0	141.165	<<< Guide value (est. value for zero magnitude error at 0Hz)
C1 set	130.450
Cv set	140.750	<<< Selected value
			x						y				| - - - Determining the gradient of Phi vs f: y=phi, x=f
			f	Rv	C1'm	C1'0	Cv	Xcv	Xcvset	U	Rbal	Xbal	|Zbal|	(|Zbal|/R0)-1	Phi				esd phi	fit	weight					Phi calc	obs-calc
		Log(f)	/MHz	/Ohms	/pF	/pF	/pF				Ohms	Ohms	Ohms	Magn. err.	/deg.					w	wx	wy	wxy	wx^2	Y calc	e	we	we^2
		6.20	1.6	1440.0	117.01	117.01	141.64	-706.0	-710.5	3.87E-07	49.759	-0.152	49.760	-0.0048	-0.175				0.0500	1	400.00	636.6	-70.0	-111.5	1013.2	0.184	-0.359	-143.476	51.463125
		6.23	1.7	1438.8	116.96	116.96	141.59	-661.2	-665.2	3.97E-07	49.759	-0.123	49.759	-0.0048	-0.142				0.0500	1	400.00	680.0	-56.8	-96.6	1156.0	0.184	-0.326	-130.378	42.496379
		6.26	1.8	1436.3	116.94	116.94	141.57	-624.6	-628.2	4.05E-07	49.748	-0.084	49.748	-0.0050	-0.097				0.0500	1	400.00	720.0	-38.6	-69.5	1296.0	0.184	-0.281	-112.298	31.527096
		6.28	1.9	1436.0	116.92	116.92	141.55	-591.8	-595.1	4.12E-07	49.747	-0.076	49.747	-0.0051	-0.087				0.0500	1	400.00	760.0	-34.8	-66.1	1444.0	0.184	-0.271	-108.573	29.470010
Calibration freq. >>>		6.30	2.0	1435.7	116.92	116.92	141.55	-562.2	-565.4	4.18E-07	49.743	-0.069	49.743	-0.0051	-0.080				0.0500	1	400.00	800.0	-31.9	-63.8	1600.0	0.185	-0.264	-105.794	27.980684
		6.34	2.2	1431.4	116.89	116.88	141.52	-511.2	-514.0	4.29E-07	49.734	-0.013	49.734	-0.0053	-0.015				0.0500	1	400.00	880.0	-6.0	-13.2	1936.0	0.185	-0.200	-80.101	16.040544
		6.38	2.4	1428.2	116.82	116.82	141.45	-468.8	-471.2	4.37E-07	49.743	0.027	49.743	-0.0051	0.032				0.0500	1	400.00	960.0	12.6	30.3	2304.0	0.186	-0.154	-61.693	9.515203
		6.41	2.6	1424.1	116.82	116.82	141.45	-432.7	-434.9	4.45E-07	49.732	0.065	49.732	-0.0054	0.075				0.0500	1	400.00	1040.0	30.1	78.3	2704.0	0.186	-0.111	-44.408	4.930240
		6.45	2.8	1420.4	116.77	116.77	141.40	-402.0	-403.8	4.51E-07	49.741	0.100	49.741	-0.0052	0.115				0.0500	1	400.00	1120.0	46.1	129.0	3136.0	0.187	-0.072	-28.695	2.058460
		6.48	3.0	1416.7	116.76	116.76	141.39	-375.2	-376.9	4.57E-07	49.740	0.127	49.741	-0.0052	0.147				0.0500	1	400.00	1200.0	58.7	176.1	3600.0	0.187	-0.041	-16.288	0.663247
		6.54	3.5	1406.0	116.66	116.66	141.29	-321.8	-323.1	4.68E-07	49.764	0.201	49.764	-0.0047	0.231				0.0500	1	400.00	1400.0	92.4	323.5	4900.0	0.189	0.042	16.882	0.712534
		6.60	4.0	1394.8	116.54	116.54	141.17	-281.8	-282.7	4.78E-07	49.797	0.263	49.798	-0.0040	0.302				0.0500	1	400.00	1600.0	120.8	483.3	6400.0	0.190	0.112	44.747	5.005633
		6.65	4.5	1383.4	116.40	116.40	141.03	-250.8	-251.3	4.86E-07	49.842	0.316	49.843	-0.0031	0.363				0.0500	1	400.00	1800.0	145.1	653.0	8100.0	0.192	0.171	68.475	11.721898
		6.70	5.0	1370.8	116.30	116.30	140.93	-225.9	-226.2	4.93E-07	49.876	0.364	49.877	-0.0025	0.419				0.0500	0	0.00	0.0	0.0	0.0	0.0	0.193	0.226	0.000	0.000000
		6.78	6.0	1348.3	116.02	116.02	140.65	-188.6	-188.5	5.06E-07	49.974	0.433	49.976	-0.0005	0.496				0.0500	1	400.00	2400.0	198.4	1190.5	14400.0	0.196	0.300	120.131	36.078359
		6.85	7.0	1331.8	115.78	115.78	140.41	-161.9	-161.5	5.16E-07	50.065	0.457	50.068	0.0014	0.523				0.0500	1	400.00	2800.0	209.2	1464.1	19600.0	0.198	0.324	129.767	42.098871
		6.90	8.0	1312.6	115.67	115.66	140.30	-141.8	-141.3	5.25E-07	50.110	0.483	50.112	0.0022	0.552				0.0500	1	400.00	3200.0	220.9	1766.8	25600.0	0.201	0.351	140.354	49.248302
		6.95	9.0	1296.7	115.46	115.46	140.09	-126.2	-125.6	5.34E-07	50.186	0.496	50.189	0.0038	0.567				0.0500	1	400.00	3600.0	226.7	2040.2	32400.0	0.204	0.363	145.088	52.625962
		7.00	10.0	1284.3	115.35	115.35	139.98	-113.7	-113.1	5.40E-07	50.232	0.494	50.234	0.0047	0.563				0.0500	1	400.00	4000.0	225.2	2252.3	40000.0	0.207	0.356	142.526	50.784250
		7.08	12.0	1265.5	115.17	115.17	139.80	-94.9	-94.2	5.50E-07	50.304	0.472	50.306	0.0061	0.538				0.0500	1	400.00	4800.0	215.2	2582.9	57600.0	0.212	0.326	130.334	42.467365
		7.15	14.0	1259.2	115.08	115.08	139.71	-81.4	-80.8	5.54E-07	50.344	0.424	50.346	0.0069	0.483				0.0500	1	400.00	5600.0	193.0	2702.1	78400.0	0.218	0.265	105.897	28.035475
		7.20	16.0	1260.6	115.16	115.16	139.79	-71.2	-70.7	5.54E-07	50.324	0.366	50.325	0.0065	0.417				0.0500	1	400.00	6400.0	166.9	2670.2	102400.0	0.223	0.194	77.571	15.043332
		7.26	18.0	1264.8	115.28	115.28	139.91	-63.2	-62.8	5.52E-07	50.283	0.315	50.284	0.0057	0.359				0.0500	1	400.00	7200.0	143.6	2584.5	129600.0	0.229	0.130	52.062	6.776177
		7.30	20.0	1282.6	115.49	115.49	140.12	-56.8	-56.5	5.44E-07	50.215	0.249	50.215	0.0043	0.284				0.0500	1	400.00	8000.0	113.5	2270.7	160000.0	0.234	0.050	19.806	0.980709
		7.34	22.0	1306.2	115.74	115.74	140.37	-51.5	-51.4	5.33E-07	50.127	0.186	50.128	0.0026	0.212				0.0500	1	400.00	8800.0	85.0	1869.2	193600.0	0.240	-0.027	-10.969	0.300790
		7.38	24.0	1338.8	115.87	115.87	140.50	-47.2	-47.1	5.20E-07	50.085	0.125	50.085	0.0017	0.143				0.0500	1	400.00	9600.0	57.0	1368.3	230400.0	0.245	-0.103	-41.126	4.228284
		7.41	26.0	1386.2	116.26	116.26	140.89	-43.4	-43.5	5.01E-07	49.947	0.055	49.948	-0.0010	0.063				0.0500	1	400.00	10400.0	25.1	653.8	270400.0	0.251	-0.188	-75.197	14.136327
Cal freq		7.45	28.0	1439.5	116.68	116.68	141.31	-40.2	-40.4	4.81E-07	49.801	-0.007	49.801	-0.0040	-0.008				0.0500	1	400.00	11200.0	-3.2	-90.4	313600.0	0.256	-0.264	-105.777	27.971743
		7.48	30.0	1487.8	117.10	117.10	141.73	-37.4	-37.7	4.64E-07	49.655	-0.052	49.655	-0.0069	-0.060				0.0500	1	400.00	12000.0	-24.1	-723.4	360000.0	0.262	-0.322	-128.868	41.517651
									Sums >>>>		28	1.12E+04	113596.6	2320.0	26054.6	2067589.2		2.26E-01	0.00E+00	645.878652
									Parameter		esd					Chi^2/n-2 >>>>>>>>>>>			24.841487
phase error vs. f									D	1.03E+10					esd of an obs of unit weight >>>			4.984
									a	0.179186	0.070778	<<< phase intercept at 0Hz
									b	0.002757	0.005209	<<< gradient [degrees / MHz]
									tp	-0.0077	0.014470	ns <<< time delay

									Phase error, obs-calc

Magnitude error vs f